## Moorings
Ref Des | Serial Number | Deployment Number | Anchor Launch Date | Anchor Launch Time | Recover Date | Latitude | Longitude | Water Depth | Cruise Number | Notes
CE01ISSP | CE01ISSP-00001 | 1 | 2014-04-17 | 22:20 |  | 44° 39.504' N | 124° 05.890' W | 25 | EA01 |  | 44.6584 | -124.0982

## Asset_Cal_Info
Ref Des | Mooring Serial Number | Deployment Number | Sensor Serial Number | Calibration Cofficient Name | Calibration Cofficient Value | Notes
CE01ISSP-SP001-02-DOSTAJ000 | CE01ISSP-00001 | 1 | 209 | CC_lat | 44.6584 | Latitude
CE01ISSP-SP001-02-DOSTAJ000 | CE01ISSP-00001 | 1 | 209 | CC_lon | -124.0982 | Longitude
CE01ISSP-SP001-02-DOSTAJ000 | CE01ISSP-00001 | 1 | 209 | CC_csv | [3.16987E-03, 1.34481E-04, 2.55659E-06, 2.31204E02, -3.73966E-01, -5.84033E01, 4.54673E00] | Requires TEMPWAT, PRESWAT and PRACSAL from CE06ISSP-SP001-09-CTDPFJ000
CE01ISSP-SP001-02-DOSTAJ000 | CE01ISSP-00001 | 1 | 209 | CC_conc_coef | [[0,1]] | 2D vector required by DPA; (1st cal, so default values for offset: 0 and slope: 1)
CE01ISSP-SP001-04-OPTAAJ000 | CE01ISSP-00001 | 1 | ACS-137 | CC_cwlngth | [399.6, 403.0, 407.1, 410.5, 413.9, 418.0, 422.2, 426.5, 430.4, 434.1, 438.4, 442.5, 446.9, 451.5, 455.9, 460.1, 464.1, 468.7, 473.4, 478.0, 482.7, 487.1, 491.4, 495.6, 500.0, 504.5, 509.2, 513.7, 518.2, 522.9, 527.2, 531.3, 535.6, 539.9, 544.0, 548.3, 552.4, 556.7, 560.8, 564.9, 568.8, 572.5, 577.3, 581.0, 584.9, 589.0, 593.0, 597.2, 601.3, 606.0, 610.0, 614.4, 618.7, 622.9, 627.0, 631.2, 635.4, 639.5, 643.8, 648.1, 652.4, 656.8, 661.1, 665.4, 669.8, 673.9, 678.1, 682.2, 686.1, 690.0, 693.9, 697.8, 701.7, 705.5, 709.2, 712.9, 716.6, 720.3, 723.8, 727.6, 731.0, 734.6, 737.8, 741.1, 744.3] | All coefficients from external file acs137_calData.txt for calibration data
CE01ISSP-SP001-04-OPTAAJ000 | CE01ISSP-00001 | 1 | ACS-137 | CC_ccwo | [0.179822, 0.268321, 0.357552, 0.426406, 0.491371, 0.553646, 0.611358, 0.656063, 0.706421, 0.753010, 0.792979, 0.833794, 0.868877, 0.903493, 0.928804, 0.956467, 0.986049, 1.010478, 1.035533, 1.058790, 1.079594, 1.098804, 1.118596, 1.134142, 1.148499, 1.163100, 1.174872, 1.186160, 1.198591, 1.210558, 1.221176, 1.231289, 1.241309, 1.251789, 1.262101, 1.272433, 1.282833, 1.291190, 1.299804, 1.306366, 1.311018, 1.311260, 1.306834, 1.299591, 1.288337, 1.271924, 1.250748, 1.231259, 1.219767, 1.216313, 1.219018, 1.223595, 1.226356, 1.229948, 1.233118, 1.235357, 1.236833, 1.236454, 1.232742, 1.225039, 1.215319, 1.206329, 1.200896, 1.199738, 1.198074, 1.194296, 1.186506, 1.171353, 1.148190, 1.115522, 1.072088, 1.016308, 0.945396, 0.854782, 0.741697, 0.602804, 0.434224, 0.237658, 0.009896, -0.233918, -0.470888, -0.667426, -0.806559, -0.892563, -0.791684]
CE01ISSP-SP001-04-OPTAAJ000 | CE01ISSP-00001 | 1 | ACS-137 | CC_tcal | 20.3
CE01ISSP-SP001-04-OPTAAJ000 | CE01ISSP-00001 | 1 | ACS-137 | CC_tbins | [3.345467, 4.467387, 5.463356, 6.462875, 7.456078, 8.467955, 9.502895, 10.518529, 11.526905, 12.483529, 13.481778, 14.492821, 15.477273, 16.471875, 17.476786, 18.478400, 19.463182, 20.500435, 21.490500, 22.493333, 23.532500, 24.509130, 25.563519, 26.471167, 27.481667, 28.495769, 29.490182, 30.488542, 31.484091, 32.488049, 33.487000, 34.505588, 35.496364, 36.486364, 37.519231, 38.277719]
CE01ISSP-SP001-04-OPTAAJ000 | CE01ISSP-00001 | 1 | ACS-137 | CC_awlngth | [400.0, 403.4, 407.3, 410.9, 414.4, 418.0, 422.2, 426.5, 430.6, 434.5, 438.4, 442.5, 446.9, 451.3, 455.7, 459.6, 463.8, 468.2, 473.1, 477.6, 482.4, 486.6, 490.8, 495.1, 499.3, 504.0, 508.3, 512.8, 517.7, 522.2, 526.3, 530.3, 534.4, 538.7, 543.2, 547.3, 551.6, 555.7, 559.9, 563.7, 567.8, 571.5, 575.4, 579.1, 582.9, 586.6, 590.5, 594.7, 598.9, 603.1, 607.5, 611.7, 616.0, 620.0, 624.2, 628.2, 632.5, 636.5, 640.7, 645.0, 649.1, 653.4, 657.7, 662.3, 666.4, 670.9, 674.9, 679.1, 683.0, 686.9, 690.8, 694.7, 698.6, 702.1, 705.9, 709.4, 713.6, 716.9, 720.6, 724.5, 727.8, 731.0, 734.5, 738.1, 741.3]
CE01ISSP-SP001-04-OPTAAJ000 | CE01ISSP-00001 | 1 | ACS-137 | CC_acwo | [-1.203530, -0.910227, -0.680714, -0.511974, -0.388067, -0.290236, -0.210387, -0.140154, -0.078221, -0.019971, 0.031966, 0.083321, 0.131855, 0.178054, 0.222949, 0.266135, 0.307730, 0.347821, 0.387471, 0.425542, 0.462535, 0.497752, 0.531535, 0.563411, 0.593687, 0.622323, 0.649312, 0.675783, 0.703294, 0.730730, 0.757586, 0.783923, 0.808800, 0.832309, 0.854311, 0.875186, 0.894651, 0.914251, 0.933026, 0.950502, 0.966319, 0.979580, 0.977900, 0.987035, 0.992680, 0.994005, 0.990693, 0.982037, 0.970928, 0.963689, 0.964909, 0.973750, 0.986004, 0.999199, 1.012098, 1.024501, 1.036668, 1.048045, 1.058163, 1.065900, 1.070011, 1.070174, 1.069012, 1.070346, 1.075649, 1.082627, 1.088433, 1.090609, 1.087587, 1.077285, 1.058061, 1.028178, 0.986388, 0.930345, 0.856421, 0.760304, 0.638846, 0.489823, 0.310908, 0.100165, -0.134875, -0.374097, -0.586355, -0.747043, -0.851335]
CE01ISSP-SP001-04-OPTAAJ000 | CE01ISSP-00001 | 1 | ACS-137 | CC_taarray | SheetRef:ACS137_CC_taarray
CE01ISSP-SP001-04-OPTAAJ000 | CE01ISSP-00001 | 1 | ACS-137 | CC_tcarray | SheetRef:ACS137_CC_tcarray
CE01ISSP-SP001-05-VELPTJ000 | CE01ISSP-00001 | 1 | AQD11259 | CC_lat | 44.6584 | Latitude
CE01ISSP-SP001-05-VELPTJ000 | CE01ISSP-00001 | 1 | AQD11259 | CC_lon | -124.0982 | Longitude
CE01ISSP-SP001-06-NUTNRJ000 | CE01ISSP-00001 | 1 | 337 | CC_lower_wavelength_limit_for_spectra_fit | 217 | Constant. wllower = Lower wavelength limit for spectra fit. From DPS: 217 nm (1-cm pathlength probe tip) or 220 nm (4-cm pathlength probe tip)
CE01ISSP-SP001-06-NUTNRJ000 | CE01ISSP-00001 | 1 | 337 | CC_upper_wavelength_limit_for_spectra_fit | 240 | Constant. wlupper = Upper wavelength limit for spectra fit. From DPS: 240 nm (1-cm pathlength probe tip) or 245 nm (4-cm pathlength probe tip)
CE01ISSP-SP001-06-NUTNRJ000 | CE01ISSP-00001 | 1 | 337 | CC_cal_temp | 19.99 | cal_temp = Calibration water temperature value
CE01ISSP-SP001-06-NUTNRJ000 | CE01ISSP-00001 | 1 | 337 | CC_wl | [190.58, 191.37, 192.15, 192.94, 193.73, 194.51, 195.30, 196.09, 196.87, 197.66, 198.45, 199.24, 200.02, 200.81, 201.60, 202.39, 203.18, 203.97, 204.76, 205.55, 206.34, 207.13, 207.92, 208.71, 209.50, 210.29, 211.09, 211.88, 212.67, 213.46, 214.26, 215.05, 215.84, 216.63, 217.43, 218.22, 219.02, 219.81, 220.61, 221.40, 222.19, 222.99, 223.79, 224.58, 225.38, 226.17, 226.97, 227.77, 228.56, 229.36, 230.16, 230.95, 231.75, 232.55, 233.35, 234.14, 234.94, 235.74, 236.54, 237.34, 238.14, 238.94, 239.74, 240.53, 241.33, 242.13, 242.93, 243.73, 244.53, 245.33, 246.14, 246.94, 247.74, 248.54, 249.34, 250.14, 250.94, 251.74, 252.55, 253.35, 254.15, 254.95, 255.75, 256.56, 257.36, 258.16, 258.97, 259.77, 260.57, 261.38, 262.18, 262.98, 263.79, 264.59, 265.39, 266.20, 267.00, 267.81, 268.61, 269.42, 270.22, 271.03, 271.83, 272.64, 273.44, 274.25, 275.05, 275.86, 276.66, 277.47, 278.27, 279.08, 279.88, 280.69, 281.50, 282.30, 283.11, 283.91, 284.72, 285.53, 286.33, 287.14, 287.95, 288.75, 289.56, 290.37, 291.17, 291.98, 292.79, 293.59, 294.40, 295.21, 296.02, 296.82, 297.63, 298.44, 299.24, 300.05, 300.86, 301.67, 302.47, 303.28, 304.09, 304.90, 305.70, 306.51, 307.32, 308.13, 308.93, 309.74, 310.55, 311.36, 312.17, 312.97, 313.78, 314.59, 315.40, 316.20, 317.01, 317.82, 318.63, 319.43, 320.24, 321.05, 321.86, 322.67, 323.47, 324.28, 325.09, 325.90, 326.70, 327.51, 328.32, 329.13, 329.93, 330.74, 331.55, 332.36, 333.16, 333.97, 334.78, 335.59, 336.39, 337.20, 338.01, 338.81, 339.62, 340.43, 341.24, 342.04, 342.85, 343.66, 344.46, 345.27, 346.08, 346.88, 347.69, 348.50, 349.30, 350.11, 350.91, 351.72, 352.53, 353.33, 354.14, 354.94, 355.75, 356.56, 357.36, 358.17, 358.97, 359.78, 360.58, 361.39, 362.19, 363.00, 363.80, 364.61, 365.41, 366.21, 367.02, 367.82, 368.63, 369.43, 370.24, 371.04, 371.84, 372.65, 373.45, 374.25, 375.06, 375.86, 376.66, 377.46, 378.27, 379.07, 379.87, 380.67, 381.48, 382.28, 383.08, 383.88, 384.68, 385.48, 386.28, 387.08, 387.89, 388.69, 389.49, 390.29, 391.09, 391.89, 392.69, 393.49, 394.29, 395.08] | wl = (1 x 256) array of wavelength bins
CE01ISSP-SP001-06-NUTNRJ000 | CE01ISSP-00001 | 1 | 337 | CC_eno3 | [-0.00534629, -0.00195085, 0.00739095, 0.00218636, -0.01877935, -0.01487355, -0.00995921, -0.00957819, 0.00211400, 0.00283378, 0.00253494, 0.00316788, 0.00613512, 0.01783688, 0.00543902, 0.00569384, 0.02133073, -0.00223603, 0.00554651, -0.00224352, 0.00197002, 0.00412564, 0.00527500, 0.00581961, 0.00672707, 0.00740365, 0.00700136, 0.00678095, 0.00658539, 0.00623560, 0.00590691, 0.00556924, 0.00519378, 0.00487035, 0.00451810, 0.00417322, 0.00383769, 0.00353066, 0.00322004, 0.00292431, 0.00263603, 0.00236783, 0.00211067, 0.00187642, 0.00165894, 0.00144268, 0.00125270, 0.00108211, 0.00092692, 0.00080048, 0.00067677, 0.00056553, 0.00048109, 0.00039671, 0.00033417, 0.00027082, 0.00022132, 0.00017508, 0.00014043, 0.00010737, 0.00007451, 0.00005906, 0.00004670, 0.00002875, 0.00002466, 0.00001226, 0.00000620, -0.00000067, -0.00000776, -0.00000916, -0.00000590, -0.00000764, -0.00000779, -0.00001741, -0.00001580, -0.00000745, -0.00001225, -0.00001572, -0.00001515, -0.00000725, -0.00000927, -0.00000460, 0.00000176, 0.00000908, 0.00001083, 0.00001443, 0.00002007, 0.00001580, 0.00001513, 0.00000878, 0.00001962, 0.00002019, 0.00002025, 0.00001338, 0.00002931, 0.00002550, 0.00003149, 0.00002782, 0.00002856, 0.00003970, 0.00002970, 0.00004209, 0.00004025, 0.00004559, 0.00003967, 0.00004447, 0.00004655, 0.00005216, 0.00005589, 0.00006664, 0.00006364, 0.00006993, 0.00006193, 0.00006369, 0.00007288, 0.00008363, 0.00008744, 0.00006025, 0.00006847, 0.00007925, 0.00008914, 0.00008182, 0.00008744, 0.00009823, 0.00009041, 0.00008985, 0.00009007, 0.00009376, 0.00010488, 0.00010771, 0.00013304, 0.00011810, 0.00013350, 0.00013299, 0.00012464, 0.00012261, 0.00013795, 0.00014428, 0.00013030, 0.00012859, 0.00012281, 0.00013909, 0.00014513, 0.00015097, 0.00014286, 0.00015057, 0.00014774, 0.00015708, 0.00015734, 0.00016186, 0.00016722, 0.00016204, 0.00016778, 0.00018088, 0.00017961, 0.00016862, 0.00017499, 0.00017953, 0.00018268, 0.00018087, 0.00018343, 0.00019139, 0.00017599, 0.00017627, 0.00018476, 0.00019542, 0.00019450, 0.00020581, 0.00022003, 0.00020295, 0.00019712, 0.00020358, 0.00020385, 0.00019639, 0.00020345, 0.00021896, 0.00022746, 0.00021637, 0.00023011, 0.00021645, 0.00023685, 0.00023503, 0.00023761, 0.00024407, 0.00022805, 0.00022477, 0.00023579, 0.00023719, 0.00026768, 0.00025649, 0.00027195, 0.00027484, 0.00026489, 0.00026238, 0.00025311, 0.00025266, 0.00024613, 0.00025112, 0.00026154, 0.00027493, 0.00026305, 0.00026761, 0.00027897, 0.00028409, 0.00028183, 0.00029914, 0.00031363, 0.00027663, 0.00026695, 0.00027802, 0.00028361, 0.00028854, 0.00028061, 0.00031304, 0.00032605, 0.00032270, 0.00031304, 0.00030940, 0.00031333, 0.00035322, 0.00033475, 0.00032788, 0.00032789, 0.00033888, 0.00035177, 0.00031966, 0.00036216, 0.00036805, 0.00034174, 0.00031028, 0.00036438, 0.00035053, 0.00036267, 0.00036168, 0.00036551, 0.00037568, 0.00034431, 0.00038197, 0.00036145, 0.00037347, 0.00035810, 0.00038371, 0.00037917, 0.00038090, 0.00038050, 0.00038780, 0.00040384, 0.00039730, 0.00040576, 0.00039112, 0.00042753, 0.00039962, 0.00041694, 0.00043307, 0.00041420, 0.00046521] | eno3 = (1 x 256) array of wavelength-dependent nitrate extinction coefficients
CE01ISSP-SP001-06-NUTNRJ000 | CE01ISSP-00001 | 1 | 337 | CC_eswa | [0.01305860, 0.00964906, 0.00803229, 0.00773347, 0.02323338, 0.01635538, 0.01010813, 0.00595118, -0.00062711, 0.00562740, 0.01110613, 0.02622002, 0.04369427, 0.04996740, 0.06515969, 0.07272178, 0.07502826, 0.08427776, 0.08011699, 0.08318401, 0.07737251, 0.07087171, 0.06165318, 0.05410329, 0.04617845, 0.03857088, 0.03232884, 0.02659763, 0.02178287, 0.01775497, 0.01437402, 0.01157746, 0.00928713, 0.00741161, 0.00589309, 0.00467891, 0.00370906, 0.00293540, 0.00231385, 0.00182382, 0.00143697, 0.00113398, 0.00089193, 0.00071024, 0.00056333, 0.00045080, 0.00035959, 0.00029797, 0.00024372, 0.00019077, 0.00015305, 0.00013678, 0.00011133, 0.00009008, 0.00008114, 0.00007133, 0.00004796, 0.00004308, 0.00003400, 0.00002595, 0.00003426, 0.00002028, 0.00001415, 0.00001605, 0.00000970, 0.00000330, -0.00000524, 0.00000532, -0.00000319, -0.00000229, -0.00000999, -0.00000752, -0.00000987, -0.00000206, -0.00000008, -0.00000985, 0.00000237, 0.00000055, -0.00000676, -0.00000264, 0.00000441, 0.00000144, -0.00000488, -0.00000014, 0.00000752, 0.00001061, 0.00000306, 0.00000017, -0.00000213, 0.00001481, -0.00000424, -0.00000161, 0.00000384, 0.00001258, -0.00000362, 0.00000267, 0.00001481, 0.00000226, 0.00000217, 0.00001242, 0.00000997, 0.00000478, 0.00000993, 0.00000373, 0.00000869, 0.00000218, 0.00000917, 0.00001200, -0.00000412, -0.00000184, 0.00000287, 0.00000387, 0.00000664, 0.00001450, 0.00000273, -0.00000025, 0.00000312, 0.00001157, 0.00000934, 0.00001350, 0.00000115, 0.00000431, 0.00001948, 0.00002057, 0.00002327, 0.00003335, 0.00004929, 0.00004997, 0.00002530, 0.00001962, 0.00000624, 0.00001927, 0.00001376, 0.00001472, 0.00003412, 0.00004296, 0.00001875, 0.00002815, 0.00004435, 0.00003892, 0.00004548, 0.00003151, 0.00003960, 0.00004667, 0.00004781, 0.00003534, 0.00004494, 0.00004468, 0.00005167, 0.00005519, 0.00005855, 0.00005775, 0.00005832, 0.00005747, 0.00005889, 0.00006687, 0.00006258, 0.00006705, 0.00007168, 0.00006533, 0.00007965, 0.00005774, 0.00007081, 0.00007665, 0.00008002, 0.00006984, 0.00007684, 0.00006858, 0.00006916, 0.00008420, 0.00008434, 0.00011238, 0.00011029, 0.00012591, 0.00010403, 0.00010553, 0.00009611, 0.00008987, 0.00008675, 0.00009395, 0.00009661, 0.00009997, 0.00010941, 0.00011355, 0.00014112, 0.00013990, 0.00013313, 0.00014122, 0.00010454, 0.00012787, 0.00012955, 0.00011490, 0.00011150, 0.00012769, 0.00014145, 0.00013132, 0.00015737, 0.00015466, 0.00015026, 0.00014071, 0.00018056, 0.00017755, 0.00016225, 0.00016260, 0.00017684, 0.00015186, 0.00014175, 0.00016650, 0.00018722, 0.00018850, 0.00019151, 0.00018487, 0.00019290, 0.00018057, 0.00014851, 0.00016983, 0.00018315, 0.00016990, 0.00018929, 0.00015949, 0.00019809, 0.00019937, 0.00019559, 0.00021248, 0.00018663, 0.00022303, 0.00018614, 0.00019402, 0.00019508, 0.00020624, 0.00016768, 0.00019848, 0.00019164, 0.00021599, 0.00020327, 0.00022334, 0.00023639, 0.00021710, 0.00025876, 0.00023415, 0.00024454, 0.00023183, 0.00022118, 0.00022510, 0.00024057, 0.00023277, 0.00024675, 0.00026366, 0.00022837, 0.00026162, 0.00022918, 0.00023026, 0.00021785, 0.00023102, 0.00025775, 0.00022232] | eswa = (1 x 256) array of seawater extinction coefficients
CE01ISSP-SP001-06-NUTNRJ000 | CE01ISSP-00001 | 1 | 337 | CC_di | [36.00000000, 47.00000000, 34.00000000, 48.00000000, 42.00000000, 32.00000000, 42.00000000, 29.00000000, 42.00000000, 66.00000000, 143.00000000, 346.00000000, 909.00000000, 2151.00000000, 4359.00000000, 7488.00000000, 11076.00000000, 14297.00000000, 16758.00000000, 18298.00000000, 19109.00000000, 19497.00000000, 19673.00000000, 19839.00000000, 20138.00000000, 20659.00000000, 21411.00000000, 22523.00000000, 23933.00000000, 25696.00000000, 27817.00000000, 30311.00000000, 33091.00000000, 36132.00000000, 39263.00000000, 42339.00000000, 45229.00000000, 47667.00000000, 49370.00000000, 50353.00000000, 50401.00000000, 49616.00000000, 48243.00000000, 46425.00000000, 44374.00000000, 42390.00000000, 40556.00000000, 39031.00000000, 37796.00000000, 36909.00000000, 36332.00000000, 36160.00000000, 36230.00000000, 36561.00000000, 37176.00000000, 37994.00000000, 38999.00000000, 40249.00000000, 41624.00000000, 43168.00000000, 44860.00000000, 46491.00000000, 48166.00000000, 49723.00000000, 51013.00000000, 51965.00000000, 52552.00000000, 52600.00000000, 52099.00000000, 51091.00000000, 49581.00000000, 47771.00000000, 45638.00000000, 43386.00000000, 41127.00000000, 38887.00000000, 36836.00000000, 34919.00000000, 33209.00000000, 31693.00000000, 30398.00000000, 29249.00000000, 28286.00000000, 27499.00000000, 26902.00000000, 26428.00000000, 26110.00000000, 25920.00000000, 25883.00000000, 25981.00000000, 26199.00000000, 26570.00000000, 27059.00000000, 27626.00000000, 28326.00000000, 29108.00000000, 29917.00000000, 30717.00000000, 31536.00000000, 32272.00000000, 32869.00000000, 33300.00000000, 33540.00000000, 33546.00000000, 33312.00000000, 32825.00000000, 32159.00000000, 31324.00000000, 30309.00000000, 29243.00000000, 28155.00000000, 27073.00000000, 26028.00000000, 25052.00000000, 24141.00000000, 23322.00000000, 22591.00000000, 21955.00000000, 21434.00000000, 20990.00000000, 20660.00000000, 20409.00000000, 20261.00000000, 20199.00000000, 20218.00000000, 20330.00000000, 20529.00000000, 20785.00000000, 21082.00000000, 21468.00000000, 21902.00000000, 22401.00000000, 22957.00000000, 23531.00000000, 24186.00000000, 24862.00000000, 25546.00000000, 26256.00000000, 27019.00000000, 27749.00000000, 28476.00000000, 29182.00000000, 29862.00000000, 30504.00000000, 31047.00000000, 31519.00000000, 31925.00000000, 32267.00000000, 32459.00000000, 32565.00000000, 32561.00000000, 32423.00000000, 32211.00000000, 31872.00000000, 31433.00000000, 30919.00000000, 30338.00000000, 29695.00000000, 29044.00000000, 28354.00000000, 27638.00000000, 26946.00000000, 26257.00000000, 25592.00000000, 24940.00000000, 24290.00000000, 23641.00000000, 23015.00000000, 22400.00000000, 21780.00000000, 21172.00000000, 20638.00000000, 20085.00000000, 19603.00000000, 19146.00000000, 18756.00000000, 18398.00000000, 18052.00000000, 17805.00000000, 17585.00000000, 17404.00000000, 17248.00000000, 17121.00000000, 16991.00000000, 16888.00000000, 16769.00000000, 16642.00000000, 16562.00000000, 16457.00000000, 16362.00000000, 16270.00000000, 16176.00000000, 16097.00000000, 16061.00000000, 16035.00000000, 16000.00000000, 16007.00000000, 15983.00000000, 15957.00000000, 15987.00000000, 16039.00000000, 16074.00000000, 16111.00000000, 16137.00000000, 16175.00000000, 16215.00000000, 16231.00000000, 16271.00000000, 16287.00000000, 16296.00000000, 16278.00000000, 16242.00000000, 16197.00000000, 16139.00000000, 16050.00000000, 15956.00000000, 15837.00000000, 15679.00000000, 15571.00000000, 15421.00000000, 15303.00000000, 15095.00000000, 14824.00000000, 14549.00000000, 14219.00000000, 13911.00000000, 13622.00000000, 13370.00000000, 13108.00000000, 12877.00000000, 12664.00000000, 12493.00000000, 12298.00000000, 12131.00000000, 12026.00000000, 11947.00000000, 11783.00000000, 11544.00000000, 11320.00000000, 11085.00000000, 10910.00000000, 10732.00000000, 10544.00000000, 10414.00000000, 10293.00000000, 10218.00000000, 10183.00000000, 10113.00000000, 9989.00000000, 9882.00000000, 9829.00000000, 9740.00000000, 9558.00000000, 9113.00000000, 8392.00000000, 7544.00000000] | di = (1 x 256) array of deionized water reference spectrum (factory Calibration suna 2013_06_21)
CE01ISSP-SP001-07-SPKIRJ000 | CE01ISSP-00001 | 1 | 237 | CC_immersion_factor | [1.368, 1.410, 1.365, 1.354, 1.372, 1.322, 1.347]
CE01ISSP-SP001-07-SPKIRJ000 | CE01ISSP-00001 | 1 | 237 | CC_offset | [2147811201.5, 2147535224.8, 2147869885.3, 2147377899.0, 2148076327.7, 2147344422.9, 2147437153.8]
CE01ISSP-SP001-07-SPKIRJ000 | CE01ISSP-00001 | 1 | 237 | CC_scale | [1.95389228132e-007, 1.89385874729e-007, 2.00792050796e-007, 2.05814024571e-007, 2.07618828172e-007, 2.10566952556e-007, 2.0654103425e-007]
CE01ISSP-SP001-08-FLORTJ000 | CE01ISSP-00001 | 1 | 1084 | CC_dark_counts_volume_scatter | 55 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
CE01ISSP-SP001-08-FLORTJ000 | CE01ISSP-00001 | 1 | 1084 | CC_scale_factor_volume_scatter | 3.08e-06 | Multiplier [m^-1 sr^-1 counts^-1]
CE01ISSP-SP001-08-FLORTJ000 | CE01ISSP-00001 | 1 | 1084 | CC_dark_counts_chlorophyll_a | 58 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
CE01ISSP-SP001-08-FLORTJ000 | CE01ISSP-00001 | 1 | 1084 | CC_scale_factor_chlorophyll_a | 0.0122 | Multiplier [ug L^-1 counts^-1]
CE01ISSP-SP001-08-FLORTJ000 | CE01ISSP-00001 | 1 | 1084 | CC_dark_counts_cdom | 48 | Measured signal output of fluormeter in clean water with black tape over the detector [counts]
CE01ISSP-SP001-08-FLORTJ000 | CE01ISSP-00001 | 1 | 1084 | CC_scale_factor_cdom | 0.0906 | Multiplier [ppb counts^-1]
CE01ISSP-SP001-08-FLORTJ000 | CE01ISSP-00001 | 1 | 1084 | CC_scattering_angle | 124 | Optical backscatter scattering angle [degrees].
CE01ISSP-SP001-08-FLORTJ000 | CE01ISSP-00001 | 1 | 1084 | CC_measurement_wavelength | 700 | Optical backscatter measurement wavelength [nm].
CE01ISSP-SP001-08-FLORTJ000 | CE01ISSP-00001 | 1 | 1084 | CC_angular_resolution | 1.076 | X (Chi) factor for high angular resolution.
CE01ISSP-SP001-08-FLORTJ000 | CE01ISSP-00001 | 1 | 1084 | CC_depolarization_ratio | 0.039 | Depolarization ratio [unitless].
CE01ISSP-SP001-09-CTDPFJ000 | CE01ISSP-00001 | 1 | 49-0308 | CC_lat | 44.6584 | Latitude
CE01ISSP-SP001-09-CTDPFJ000 | CE01ISSP-00001 | 1 | 49-0308 | CC_lon | -124.0982 | Longitude
CE01ISSP-SP001-10-PARADJ000 | CE01ISSP-00001 | 1 | 365 | CC_a0 | 4381
CE01ISSP-SP001-10-PARADJ000 | CE01ISSP-00001 | 1 | 365 | CC_a1 | 2904
CE01ISSP-SP001-10-PARADJ000 | CE01ISSP-00001 | 1 | 365 | CC_Im | 1.3589
CE01ISSP-SP001-00-SPPENG000 | CE01ISSP-00001 | 1 | WLP-001

## ACS137_CC_taarray
0.046008 | 0.038616 | 0.034189 | 0.030633 | 0.027713 | 0.024801 | 0.022116 | 0.019702 | 0.017093 | 0.015481 | 0.015147 | 0.01252 | 0.008928 | 0.009577 | 0.007696 | 0.007155 | 0.005752 | 0.005757 | 0.004236 | 0.002105 | 0.000767 | 0.001314 | 0 | -0.003164 | -0.00289 | -0.004972 | -0.005773 | -0.0052 | -0.00454 | -0.00453 | -0.005466 | -0.006406 | -0.008507 | -0.0092 | -0.011614 | -0.01137
0.042537 | 0.036946 | 0.03306 | 0.029801 | 0.027492 | 0.025002 | 0.022203 | 0.02048 | 0.018034 | 0.016726 | 0.015914 | 0.013964 | 0.011304 | 0.010473 | 0.010112 | 0.008744 | 0.007464 | 0.005362 | 0.003948 | 0.001946 | 0.001172 | 0.000262 | 0 | -0.00248 | -0.002043 | -0.004583 | -0.004992 | -0.004635 | -0.004805 | -0.004368 | -0.006723 | -0.008813 | -0.010446 | -0.010831 | -0.014073 | -0.01358
0.043115 | 0.038416 | 0.035336 | 0.032489 | 0.030057 | 0.027135 | 0.024997 | 0.022555 | 0.020393 | 0.019228 | 0.017791 | 0.015667 | 0.013552 | 0.013029 | 0.011394 | 0.010092 | 0.008747 | 0.007032 | 0.00537 | 0.002942 | 0.001764 | 0.001265 | 0 | -0.00225 | -0.002327 | -0.004171 | -0.004781 | -0.005138 | -0.006065 | -0.00618 | -0.008225 | -0.009586 | -0.011717 | -0.012542 | -0.015809 | -0.015839
0.043878 | 0.039544 | 0.036954 | 0.034264 | 0.031433 | 0.028793 | 0.026784 | 0.023817 | 0.022177 | 0.020809 | 0.019306 | 0.01721 | 0.01505 | 0.014461 | 0.012233 | 0.010651 | 0.009833 | 0.007581 | 0.005108 | 0.002744 | 0.001781 | 0.000955 | 0 | -0.002318 | -0.002986 | -0.004811 | -0.005572 | -0.006023 | -0.007244 | -0.007514 | -0.009532 | -0.011754 | -0.014287 | -0.015186 | -0.018808 | -0.018959
0.043221 | 0.039518 | 0.036684 | 0.034183 | 0.03171 | 0.028892 | 0.026332 | 0.024063 | 0.022705 | 0.021051 | 0.019975 | 0.017381 | 0.016355 | 0.014593 | 0.013038 | 0.011016 | 0.009309 | 0.007655 | 0.005453 | 0.002885 | 0.002121 | 0.000702 | 0 | -0.002386 | -0.003304 | -0.00499 | -0.006084 | -0.006539 | -0.007517 | -0.008378 | -0.010583 | -0.012643 | -0.015208 | -0.016499 | -0.020316 | -0.020358
0.041511 | 0.038191 | 0.035817 | 0.033419 | 0.030901 | 0.02826 | 0.02634 | 0.023815 | 0.02206 | 0.020759 | 0.01963 | 0.017404 | 0.015678 | 0.014926 | 0.012851 | 0.010982 | 0.009436 | 0.00781 | 0.00523 | 0.002876 | 0.002447 | 0.001 | 0 | -0.001817 | -0.003196 | -0.004607 | -0.005684 | -0.00597 | -0.007093 | -0.008098 | -0.010008 | -0.01219 | -0.01478 | -0.01601 | -0.019111 | -0.019682
0.037417 | 0.034269 | 0.032527 | 0.030387 | 0.028083 | 0.025781 | 0.02401 | 0.021743 | 0.020421 | 0.019 | 0.017739 | 0.016002 | 0.014152 | 0.013551 | 0.01176 | 0.010372 | 0.008882 | 0.007288 | 0.005062 | 0.002737 | 0.002 | 0.000832 | 0 | -0.001774 | -0.002436 | -0.004129 | -0.004938 | -0.005171 | -0.006317 | -0.006974 | -0.009036 | -0.010991 | -0.01359 | -0.01498 | -0.017872 | -0.018426
0.034866 | 0.03179 | 0.029563 | 0.02768 | 0.025457 | 0.023346 | 0.021735 | 0.020139 | 0.018407 | 0.017358 | 0.016297 | 0.014795 | 0.013411 | 0.012652 | 0.011058 | 0.0093 | 0.008381 | 0.006867 | 0.004741 | 0.002691 | 0.001716 | 0.000637 | 0 | -0.001355 | -0.002443 | -0.003704 | -0.004677 | -0.004526 | -0.005612 | -0.006404 | -0.008255 | -0.01001 | -0.012012 | -0.013492 | -0.016152 | -0.016534
0.0312 | 0.028177 | 0.026195 | 0.02433 | 0.02227 | 0.020543 | 0.019072 | 0.017102 | 0.015948 | 0.015189 | 0.014157 | 0.012954 | 0.01166 | 0.010825 | 0.009675 | 0.008529 | 0.007222 | 0.005664 | 0.004205 | 0.002227 | 0.001055 | 0.000415 | 0 | -0.001326 | -0.002234 | -0.003411 | -0.003867 | -0.00402 | -0.004773 | -0.005597 | -0.006942 | -0.008896 | -0.011072 | -0.012037 | -0.01447 | -0.014979
0.027693 | 0.024791 | 0.023094 | 0.021474 | 0.019464 | 0.017743 | 0.016514 | 0.015014 | 0.013888 | 0.013231 | 0.012383 | 0.011222 | 0.010159 | 0.009616 | 0.008586 | 0.00719 | 0.006435 | 0.005223 | 0.003644 | 0.001914 | 0.001161 | 0.000443 | 0 | -0.00127 | -0.00217 | -0.002973 | -0.003355 | -0.0034 | -0.003972 | -0.004987 | -0.006092 | -0.00759 | -0.009722 | -0.010993 | -0.013021 | -0.0136
0.024787 | 0.022258 | 0.020647 | 0.019176 | 0.01742 | 0.015684 | 0.014677 | 0.013243 | 0.01208 | 0.011556 | 0.010946 | 0.009946 | 0.009138 | 0.008737 | 0.007766 | 0.006708 | 0.005749 | 0.004525 | 0.003223 | 0.001591 | 0.000825 | 0.000446 | 0 | -0.001115 | -0.001928 | -0.002527 | -0.002931 | -0.002911 | -0.003497 | -0.004243 | -0.005478 | -0.007028 | -0.008711 | -0.009666 | -0.011774 | -0.012409
0.021585 | 0.01939 | 0.01804 | 0.016783 | 0.015198 | 0.013819 | 0.012806 | 0.011385 | 0.010436 | 0.01003 | 0.009637 | 0.00875 | 0.007953 | 0.007489 | 0.006886 | 0.005919 | 0.005196 | 0.004211 | 0.00285 | 0.001583 | 0.000872 | 0.000262 | 0 | -0.000898 | -0.001568 | -0.002209 | -0.002438 | -0.002361 | -0.003038 | -0.003686 | -0.004585 | -0.006224 | -0.007827 | -0.00886 | -0.010829 | -0.011106
0.019363 | 0.017408 | 0.016234 | 0.015135 | 0.01369 | 0.012418 | 0.011536 | 0.010294 | 0.009543 | 0.00896 | 0.008634 | 0.007892 | 0.007134 | 0.006682 | 0.006158 | 0.005174 | 0.004563 | 0.003796 | 0.002592 | 0.001428 | 0.000706 | 0.000228 | 0 | -0.000795 | -0.001294 | -0.001816 | -0.002237 | -0.002205 | -0.002846 | -0.003452 | -0.004628 | -0.005848 | -0.007552 | -0.008374 | -0.010169 | -0.01066
0.017412 | 0.015691 | 0.014759 | 0.013764 | 0.012461 | 0.011293 | 0.010309 | 0.009391 | 0.008448 | 0.008114 | 0.007953 | 0.007106 | 0.006573 | 0.006196 | 0.005581 | 0.004678 | 0.004151 | 0.003395 | 0.002495 | 0.001067 | 0.000674 | 3.9e-05 | 0 | -0.000857 | -0.001501 | -0.001834 | -0.002084 | -0.001958 | -0.002642 | -0.003376 | -0.004491 | -0.005671 | -0.007114 | -0.008 | -0.009561 | -0.010091
0.01574 | 0.014302 | 0.01358 | 0.012714 | 0.011537 | 0.01044 | 0.009666 | 0.008624 | 0.007842 | 0.007636 | 0.007369 | 0.0066 | 0.006169 | 0.005925 | 0.005449 | 0.004572 | 0.004024 | 0.003217 | 0.002289 | 0.001143 | 0.000838 | 5.9e-05 | 0 | -0.000573 | -0.001096 | -0.001485 | -0.00172 | -0.001808 | -0.002392 | -0.003167 | -0.004215 | -0.005283 | -0.006654 | -0.007416 | -0.008966 | -0.009452
0.014527 | 0.01326 | 0.012629 | 0.011877 | 0.010791 | 0.009614 | 0.008946 | 0.008009 | 0.007438 | 0.007028 | 0.006935 | 0.006311 | 0.00575 | 0.005497 | 0.004951 | 0.004274 | 0.003816 | 0.003087 | 0.002192 | 0.000976 | 0.000737 | 0.000213 | 0 | -0.000583 | -0.001186 | -0.001511 | -0.001662 | -0.001691 | -0.002627 | -0.003366 | -0.004308 | -0.005396 | -0.006692 | -0.007434 | -0.008811 | -0.009214
0.01379 | 0.012703 | 0.012086 | 0.011332 | 0.010379 | 0.009374 | 0.008555 | 0.0077 | 0.007201 | 0.006862 | 0.006828 | 0.006161 | 0.005639 | 0.005414 | 0.004984 | 0.004048 | 0.003721 | 0.003082 | 0.002118 | 0.001101 | 0.000759 | 0.000216 | 0 | -0.000597 | -0.001184 | -0.001518 | -0.001751 | -0.001947 | -0.002754 | -0.003539 | -0.004422 | -0.005565 | -0.006766 | -0.007537 | -0.008837 | -0.009163
0.012543 | 0.011574 | 0.011144 | 0.010525 | 0.009549 | 0.008673 | 0.00803 | 0.007184 | 0.006723 | 0.006482 | 0.006281 | 0.005854 | 0.005194 | 0.004992 | 0.00458 | 0.003814 | 0.003451 | 0.002743 | 0.001959 | 0.000935 | 0.000687 | 0.000105 | 0 | -0.000692 | -0.001173 | -0.001591 | -0.001863 | -0.002125 | -0.002902 | -0.003845 | -0.004739 | -0.005723 | -0.006918 | -0.007474 | -0.008805 | -0.009164
0.01198 | 0.011028 | 0.010619 | 0.010011 | 0.009109 | 0.008293 | 0.007609 | 0.006829 | 0.00636 | 0.0062 | 0.00614 | 0.005654 | 0.005143 | 0.004862 | 0.004394 | 0.003629 | 0.003294 | 0.002689 | 0.001905 | 0.000907 | 0.000666 | 0.00013 | 0 | -0.000527 | -0.001178 | -0.001518 | -0.001782 | -0.002193 | -0.003075 | -0.004039 | -0.00491 | -0.005911 | -0.006964 | -0.00759 | -0.008803 | -0.009256
0.01141 | 0.01055 | 0.010146 | 0.009571 | 0.008772 | 0.007999 | 0.007373 | 0.006629 | 0.006245 | 0.006094 | 0.00599 | 0.005529 | 0.00495 | 0.004744 | 0.004335 | 0.003507 | 0.003113 | 0.002569 | 0.001904 | 0.000959 | 0.000698 | 0.000242 | 0 | -0.000572 | -0.001179 | -0.001574 | -0.00193 | -0.002426 | -0.003291 | -0.004351 | -0.005273 | -0.006235 | -0.007253 | -0.007898 | -0.008927 | -0.009331
0.01087 | 0.010027 | 0.009676 | 0.00918 | 0.008413 | 0.007705 | 0.007076 | 0.00639 | 0.005991 | 0.005909 | 0.005841 | 0.005415 | 0.004911 | 0.004576 | 0.00425 | 0.003459 | 0.003009 | 0.002547 | 0.001877 | 0.00097 | 0.000661 | 0.000233 | 0 | -0.000508 | -0.001209 | -0.001644 | -0.002019 | -0.002626 | -0.003583 | -0.004605 | -0.005614 | -0.006502 | -0.007517 | -0.008026 | -0.009049 | -0.009434
0.010403 | 0.009618 | 0.009267 | 0.008818 | 0.008182 | 0.007515 | 0.006972 | 0.006306 | 0.005913 | 0.005771 | 0.005722 | 0.005371 | 0.004799 | 0.004479 | 0.004027 | 0.003438 | 0.002982 | 0.002523 | 0.00185 | 0.001072 | 0.00071 | 0.000335 | 0 | -0.000589 | -0.001241 | -0.001699 | -0.002153 | -0.002916 | -0.004066 | -0.005083 | -0.006069 | -0.006993 | -0.007895 | -0.00852 | -0.009424 | -0.00967
0.01018 | 0.009398 | 0.009084 | 0.008655 | 0.008069 | 0.007418 | 0.006818 | 0.006126 | 0.005794 | 0.005744 | 0.005667 | 0.005292 | 0.004723 | 0.00443 | 0.003994 | 0.0033 | 0.002896 | 0.002409 | 0.001769 | 0.00099 | 0.000674 | 0.000234 | 0 | -0.000674 | -0.001285 | -0.001786 | -0.002319 | -0.003244 | -0.004354 | -0.0055 | -0.00646 | -0.007325 | -0.008186 | -0.008741 | -0.009684 | -0.00996
0.009588 | 0.008909 | 0.008627 | 0.008287 | 0.00774 | 0.007126 | 0.006552 | 0.005927 | 0.005578 | 0.0055 | 0.005499 | 0.005064 | 0.004562 | 0.004214 | 0.003825 | 0.003159 | 0.002808 | 0.002341 | 0.001725 | 0.001035 | 0.000725 | 0.000325 | 0 | -0.000595 | -0.001287 | -0.001756 | -0.002334 | -0.00328 | -0.004468 | -0.005653 | -0.006605 | -0.007448 | -0.008286 | -0.008846 | -0.009583 | -0.009919
0.009384 | 0.008704 | 0.008438 | 0.00813 | 0.007539 | 0.007009 | 0.006432 | 0.005796 | 0.005511 | 0.005399 | 0.005272 | 0.004901 | 0.004381 | 0.00406 | 0.003655 | 0.003024 | 0.002629 | 0.00221 | 0.001679 | 0.00102 | 0.000628 | 0.000243 | 0 | -0.000666 | -0.001294 | -0.001828 | -0.002439 | -0.003481 | -0.004739 | -0.005889 | -0.006905 | -0.007709 | -0.008475 | -0.009046 | -0.009783 | -0.010032
0.009025 | 0.008423 | 0.008153 | 0.007814 | 0.007376 | 0.006783 | 0.00626 | 0.005577 | 0.00542 | 0.005217 | 0.005156 | 0.004779 | 0.004293 | 0.003957 | 0.003551 | 0.0029 | 0.002579 | 0.002087 | 0.001695 | 0.00099 | 0.000637 | 0.000301 | 0 | -0.000599 | -0.001262 | -0.00174 | -0.002311 | -0.003456 | -0.00477 | -0.005866 | -0.006962 | -0.00772 | -0.00848 | -0.008957 | -0.009583 | -0.009897
0.008554 | 0.007989 | 0.00775 | 0.007456 | 0.006977 | 0.006494 | 0.005955 | 0.005341 | 0.005054 | 0.004937 | 0.004846 | 0.004462 | 0.003998 | 0.003695 | 0.003349 | 0.002746 | 0.002439 | 0.002039 | 0.001589 | 0.000828 | 0.000567 | 0.000256 | 0 | -0.000555 | -0.001177 | -0.001597 | -0.002226 | -0.003298 | -0.004639 | -0.005762 | -0.006782 | -0.007569 | -0.00827 | -0.008681 | -0.009307 | -0.009595
0.00832 | 0.007748 | 0.007496 | 0.007161 | 0.006745 | 0.006232 | 0.005666 | 0.00511 | 0.004833 | 0.004693 | 0.00457 | 0.004208 | 0.003721 | 0.003499 | 0.003025 | 0.002491 | 0.002262 | 0.001815 | 0.001388 | 0.000831 | 0.00061 | 0.000285 | 0 | -0.000552 | -0.001051 | -0.001494 | -0.002084 | -0.003199 | -0.004537 | -0.005626 | -0.00659 | -0.007341 | -0.007975 | -0.008351 | -0.008977 | -0.009245
0.007946 | 0.007378 | 0.007133 | 0.00683 | 0.00637 | 0.005849 | 0.005377 | 0.004842 | 0.004533 | 0.004369 | 0.004236 | 0.003946 | 0.003481 | 0.003232 | 0.00291 | 0.002373 | 0.002079 | 0.001742 | 0.001288 | 0.000723 | 0.000531 | 0.000312 | 0 | -0.000487 | -0.001021 | -0.00147 | -0.002028 | -0.003119 | -0.004389 | -0.005493 | -0.006438 | -0.007091 | -0.007734 | -0.008077 | -0.008647 | -0.008911
0.007576 | 0.006989 | 0.006765 | 0.006491 | 0.005999 | 0.005483 | 0.005041 | 0.00444 | 0.004187 | 0.004039 | 0.003888 | 0.0036 | 0.003229 | 0.002959 | 0.002575 | 0.002141 | 0.001908 | 0.00159 | 0.001249 | 0.000657 | 0.000472 | 0.000214 | 0 | -0.000447 | -0.000868 | -0.001279 | -0.001852 | -0.002889 | -0.004096 | -0.005097 | -0.006013 | -0.006623 | -0.007168 | -0.007525 | -0.00808 | -0.008286
0.007564 | 0.006952 | 0.006628 | 0.006312 | 0.005851 | 0.005346 | 0.00484 | 0.004223 | 0.003999 | 0.003805 | 0.003635 | 0.003368 | 0.002944 | 0.002728 | 0.002408 | 0.001966 | 0.001664 | 0.001428 | 0.001076 | 0.000599 | 0.000443 | 0.000198 | 0 | -0.000462 | -0.000913 | -0.001242 | -0.001763 | -0.002773 | -0.003921 | -0.004841 | -0.005715 | -0.006321 | -0.006825 | -0.00717 | -0.00769 | -0.007848
0.007257 | 0.006623 | 0.006335 | 0.006029 | 0.00554 | 0.005019 | 0.00457 | 0.004022 | 0.003688 | 0.003543 | 0.003395 | 0.003111 | 0.002732 | 0.002532 | 0.002184 | 0.001856 | 0.001599 | 0.001355 | 0.001036 | 0.000557 | 0.000415 | 0.000208 | 0 | -0.000373 | -0.000799 | -0.001124 | -0.001627 | -0.002504 | -0.003588 | -0.004483 | -0.00522 | -0.005818 | -0.00629 | -0.006599 | -0.007086 | -0.007273
0.007121 | 0.006443 | 0.006148 | 0.005826 | 0.00535 | 0.00485 | 0.004407 | 0.003819 | 0.003492 | 0.003336 | 0.003163 | 0.002922 | 0.002543 | 0.002373 | 0.001999 | 0.001703 | 0.001417 | 0.001175 | 0.000918 | 0.000504 | 0.000302 | 0.000184 | 0 | -0.000381 | -0.000785 | -0.001098 | -0.00151 | -0.002314 | -0.003289 | -0.004053 | -0.004797 | -0.005327 | -0.005773 | -0.00605 | -0.006499 | -0.006673
0.007101 | 0.006379 | 0.006013 | 0.00569 | 0.005225 | 0.004703 | 0.004252 | 0.003698 | 0.003376 | 0.003213 | 0.003013 | 0.002717 | 0.002412 | 0.00222 | 0.001867 | 0.001579 | 0.001281 | 0.001096 | 0.000854 | 0.000463 | 0.000356 | 0.000163 | 0 | -0.000385 | -0.000799 | -0.001092 | -0.001463 | -0.002125 | -0.002988 | -0.003687 | -0.004415 | -0.004852 | -0.005262 | -0.005561 | -0.005972 | -0.006082
0.006969 | 0.006252 | 0.005898 | 0.005572 | 0.005101 | 0.004571 | 0.004166 | 0.003589 | 0.00328 | 0.003078 | 0.002878 | 0.002632 | 0.002269 | 0.002081 | 0.001748 | 0.001471 | 0.00127 | 0.001052 | 0.00078 | 0.000409 | 0.000288 | 0.000103 | 0 | -0.000371 | -0.000782 | -0.001021 | -0.001343 | -0.001922 | -0.002726 | -0.003285 | -0.003948 | -0.004373 | -0.004782 | -0.005002 | -0.005395 | -0.005555
0.007158 | 0.006411 | 0.00604 | 0.00566 | 0.005208 | 0.004699 | 0.004216 | 0.003659 | 0.00337 | 0.003136 | 0.002933 | 0.002617 | 0.00227 | 0.002105 | 0.001771 | 0.00143 | 0.001223 | 0.00107 | 0.000791 | 0.000421 | 0.00025 | 0.000136 | 0 | -0.000368 | -0.000753 | -0.000999 | -0.00131 | -0.001773 | -0.002459 | -0.002984 | -0.003572 | -0.003966 | -0.004313 | -0.00452 | -0.004923 | -0.005073
0.007047 | 0.006299 | 0.005912 | 0.005579 | 0.005126 | 0.004654 | 0.004141 | 0.003589 | 0.003255 | 0.003082 | 0.002845 | 0.002535 | 0.002233 | 0.001987 | 0.001688 | 0.001381 | 0.001186 | 0.001029 | 0.000784 | 0.000384 | 0.000249 | 0.000115 | 0 | -0.000358 | -0.000728 | -0.00099 | -0.001228 | -0.001616 | -0.002197 | -0.002649 | -0.00318 | -0.003555 | -0.003879 | -0.004102 | -0.004492 | -0.004617
0.007283 | 0.006549 | 0.006159 | 0.005802 | 0.005327 | 0.004841 | 0.004356 | 0.003798 | 0.003468 | 0.00324 | 0.002973 | 0.002625 | 0.002322 | 0.002035 | 0.001765 | 0.001439 | 0.001189 | 0.00105 | 0.000791 | 0.000399 | 0.000307 | 0.000113 | 0 | -0.00033 | -0.000722 | -0.000969 | -0.001224 | -0.001552 | -0.002062 | -0.002497 | -0.002957 | -0.003266 | -0.003613 | -0.003802 | -0.004163 | -0.004322
0.007023 | 0.006333 | 0.005994 | 0.00565 | 0.005227 | 0.004716 | 0.00427 | 0.003718 | 0.003409 | 0.003178 | 0.002899 | 0.002583 | 0.002295 | 0.001998 | 0.00174 | 0.001411 | 0.001192 | 0.001036 | 0.000793 | 0.0004 | 0.000278 | 0.000101 | 0 | -0.000364 | -0.00078 | -0.000969 | -0.001226 | -0.0015 | -0.001954 | -0.002313 | -0.002749 | -0.003062 | -0.003378 | -0.003541 | -0.003902 | -0.004029
0.006609 | 0.005984 | 0.005668 | 0.005387 | 0.004953 | 0.004504 | 0.004091 | 0.003566 | 0.00327 | 0.003057 | 0.002816 | 0.00251 | 0.00218 | 0.001975 | 0.001696 | 0.001383 | 0.001171 | 0.001003 | 0.000678 | 0.000391 | 0.000262 | 7.3e-05 | 0 | -0.000355 | -0.00075 | -0.000954 | -0.001153 | -0.001409 | -0.001791 | -0.002114 | -0.002541 | -0.002823 | -0.003138 | -0.003337 | -0.003661 | -0.003791
0.006444 | 0.005873 | 0.005609 | 0.005328 | 0.004906 | 0.004447 | 0.00405 | 0.003556 | 0.00322 | 0.003015 | 0.00277 | 0.002502 | 0.002144 | 0.001965 | 0.001688 | 0.001369 | 0.001206 | 0.001019 | 0.000741 | 0.000358 | 0.000265 | 0.000114 | 0 | -0.000355 | -0.000725 | -0.000925 | -0.001144 | -0.001355 | -0.001679 | -0.001999 | -0.002383 | -0.002669 | -0.002988 | -0.003147 | -0.003506 | -0.003599
0.005917 | 0.005383 | 0.005139 | 0.004933 | 0.004511 | 0.004107 | 0.003737 | 0.003277 | 0.002987 | 0.002822 | 0.002634 | 0.002373 | 0.002063 | 0.001898 | 0.001618 | 0.001369 | 0.001147 | 0.000924 | 0.000686 | 0.000387 | 0.00027 | 0.000125 | 0 | -0.000326 | -0.00068 | -0.000876 | -0.001058 | -0.001251 | -0.001574 | -0.001856 | -0.002232 | -0.00252 | -0.002785 | -0.002966 | -0.003297 | -0.003417
0.001909 | 0.001715 | 0.001664 | 0.001463 | 0.001119 | 0.000735 | 0.000375 | 8e-06 | -0.000167 | -0.000197 | -0.000189 | -0.000198 | -0.000221 | -0.00014 | -0.000112 | -0.000103 | -0.000104 | -6.8e-05 | -0.00014 | -0.000267 | -0.000241 | -0.000208 | 0 | 0.000151 | 0.000272 | 0.000519 | 0.000804 | 0.00103 | 0.001049 | 0.000996 | 0.000826 | 0.000574 | 0.000302 | 0.000105 | -0.000244 | -0.0004
0.001421 | 0.00125 | 0.001248 | 0.001098 | 0.000786 | 0.000437 | 0.000118 | -0.000257 | -0.000426 | -0.000425 | -0.000397 | -0.000384 | -0.000404 | -0.00029 | -0.000222 | -0.000232 | -0.000184 | -0.000118 | -0.000175 | -0.000299 | -0.00025 | -0.000235 | 0 | 0.000206 | 0.00033 | 0.000601 | 0.000908 | 0.001174 | 0.001145 | 0.001144 | 0.000959 | 0.000711 | 0.000476 | 0.000284 | -5.3e-05 | -0.000197
0.00137 | 0.001241 | 0.00122 | 0.001025 | 0.000682 | 0.000345 | 1.2e-05 | -0.000348 | -0.000562 | -0.00055 | -0.000517 | -0.000499 | -0.000488 | -0.000376 | -0.000356 | -0.000304 | -0.000258 | -0.000202 | -0.000244 | -0.000341 | -0.000262 | -0.000247 | 0 | 0.000237 | 0.000414 | 0.000707 | 0.001062 | 0.00131 | 0.001389 | 0.001387 | 0.001188 | 0.000982 | 0.000743 | 0.000531 | 0.000227 | 6.6e-05
0.000976 | 0.000824 | 0.000801 | 0.000623 | 0.000308 | -5e-05 | -0.000344 | -0.000684 | -0.000888 | -0.000869 | -0.000797 | -0.000797 | -0.000751 | -0.000607 | -0.000524 | -0.000479 | -0.000423 | -0.000293 | -0.00031 | -0.000408 | -0.000326 | -0.000297 | 0 | 0.000271 | 0.000486 | 0.000836 | 0.001224 | 0.001509 | 0.001601 | 0.001587 | 0.001459 | 0.001229 | 0.000999 | 0.000815 | 0.00049 | 0.000345
0.000559 | 0.000383 | 0.000343 | 0.000189 | -0.000118 | -0.000416 | -0.000749 | -0.001037 | -0.001191 | -0.00116 | -0.001043 | -0.000995 | -0.00094 | -0.00081 | -0.000677 | -0.000609 | -0.00051 | -0.000404 | -0.000427 | -0.000493 | -0.000369 | -0.000304 | 0 | 0.0003 | 0.000553 | 0.000933 | 0.001348 | 0.001663 | 0.001768 | 0.001814 | 0.001687 | 0.001479 | 0.001255 | 0.001084 | 0.000766 | 0.000648
-1.1e-05 | -0.000217 | -0.000274 | -0.000398 | -0.000644 | -0.000931 | -0.001178 | -0.001483 | -0.001611 | -0.001533 | -0.001381 | -0.001289 | -0.001213 | -0.001025 | -0.000844 | -0.000759 | -0.00063 | -0.000493 | -0.000455 | -0.000551 | -0.000395 | -0.000322 | 0 | 0.000353 | 0.000658 | 0.001081 | 0.001528 | 0.001912 | 0.00206 | 0.002098 | 0.001996 | 0.001838 | 0.001594 | 0.001441 | 0.001129 | 0.001002
-0.000562 | -0.000829 | -0.00094 | -0.001052 | -0.001261 | -0.001501 | -0.001719 | -0.001983 | -0.002057 | -0.001941 | -0.001743 | -0.001603 | -0.001457 | -0.001272 | -0.001069 | -0.000982 | -0.000785 | -0.000632 | -0.000581 | -0.000613 | -0.000438 | -0.000385 | 0 | 0.000389 | 0.000728 | 0.001209 | 0.001731 | 0.002125 | 0.00229 | 0.002376 | 0.002297 | 0.002117 | 0.001923 | 0.00177 | 0.00149 | 0.00135
-0.001339 | -0.00164 | -0.001742 | -0.001765 | -0.001884 | -0.002078 | -0.002222 | -0.002388 | -0.002402 | -0.002223 | -0.001952 | -0.001757 | -0.001632 | -0.00139 | -0.001152 | -0.001018 | -0.000832 | -0.00066 | -0.000622 | -0.000622 | -0.000466 | -0.000373 | 0 | 0.000421 | 0.000793 | 0.001298 | 0.001807 | 0.002256 | 0.002438 | 0.002491 | 0.002448 | 0.002294 | 0.002096 | 0.001977 | 0.001675 | 0.001538
-0.00207 | -0.002399 | -0.002499 | -0.002449 | -0.002505 | -0.002584 | -0.002675 | -0.002837 | -0.002799 | -0.00255 | -0.002241 | -0.002016 | -0.001819 | -0.001558 | -0.001284 | -0.001169 | -0.000917 | -0.000739 | -0.000672 | -0.000667 | -0.000474 | -0.000396 | 0 | 0.000455 | 0.000855 | 0.001371 | 0.001922 | 0.002354 | 0.002563 | 0.002668 | 0.002632 | 0.002455 | 0.002283 | 0.002159 | 0.001876 | 0.001722
-0.002702 | -0.003046 | -0.003137 | -0.003022 | -0.002994 | -0.00302 | -0.003077 | -0.003148 | -0.003057 | -0.002747 | -0.002405 | -0.002109 | -0.001895 | -0.001586 | -0.001302 | -0.001158 | -0.000966 | -0.000764 | -0.000662 | -0.000718 | -0.000489 | -0.000405 | 0 | 0.000435 | 0.000832 | 0.001335 | 0.001917 | 0.002369 | 0.002554 | 0.00263 | 0.002581 | 0.002453 | 0.002256 | 0.002127 | 0.001837 | 0.001714
-0.002864 | -0.003218 | -0.00332 | -0.00318 | -0.003154 | -0.00316 | -0.003161 | -0.003231 | -0.003107 | -0.002789 | -0.00239 | -0.002088 | -0.001886 | -0.001585 | -0.001309 | -0.001169 | -0.000936 | -0.000732 | -0.000676 | -0.000668 | -0.000446 | -0.000358 | 0 | 0.000439 | 0.000827 | 0.001347 | 0.001896 | 0.002379 | 0.002535 | 0.002629 | 0.002559 | 0.002429 | 0.002234 | 0.002148 | 0.00185 | 0.001728
-0.002951 | -0.003248 | -0.003294 | -0.003087 | -0.002991 | -0.002975 | -0.002979 | -0.003049 | -0.00293 | -0.00264 | -0.00224 | -0.001971 | -0.001777 | -0.001485 | -0.001224 | -0.001086 | -0.000878 | -0.000706 | -0.000636 | -0.000617 | -0.00046 | -0.000379 | 0 | 0.000404 | 0.000763 | 0.001245 | 0.001775 | 0.002199 | 0.002363 | 0.002437 | 0.002369 | 0.002256 | 0.002046 | 0.001947 | 0.00168 | 0.00153
-0.002822 | -0.00308 | -0.003108 | -0.002917 | -0.002807 | -0.002776 | -0.002788 | -0.002831 | -0.002706 | -0.002373 | -0.001979 | -0.001673 | -0.001468 | -0.00121 | -0.001 | -0.000895 | -0.000725 | -0.000571 | -0.00052 | -0.000529 | -0.000426 | -0.000318 | 0 | 0.000372 | 0.000723 | 0.001153 | 0.001681 | 0.002073 | 0.002215 | 0.002258 | 0.002185 | 0.002062 | 0.001879 | 0.001777 | 0.001507 | 0.001378
-0.002755 | -0.002972 | -0.002972 | -0.002762 | -0.002692 | -0.002691 | -0.002706 | -0.002761 | -0.002659 | -0.002329 | -0.001971 | -0.001686 | -0.001504 | -0.001235 | -0.001042 | -0.000907 | -0.000726 | -0.000582 | -0.000514 | -0.000555 | -0.000427 | -0.000286 | 0 | 0.000329 | 0.000599 | 0.000963 | 0.001431 | 0.00177 | 0.001919 | 0.001951 | 0.001873 | 0.001748 | 0.001575 | 0.00145 | 0.001198 | 0.001064
-0.002514 | -0.002704 | -0.002688 | -0.002483 | -0.002417 | -0.002414 | -0.002436 | -0.002539 | -0.002445 | -0.002136 | -0.001765 | -0.001537 | -0.00136 | -0.001116 | -0.000901 | -0.000817 | -0.000663 | -0.000511 | -0.000453 | -0.00049 | -0.000391 | -0.00029 | 0 | 0.000324 | 0.00062 | 0.001015 | 0.001459 | 0.00172 | 0.001754 | 0.001794 | 0.001711 | 0.001565 | 0.001366 | 0.001288 | 0.001027 | 0.000936
-0.002244 | -0.002404 | -0.002385 | -0.002187 | -0.00215 | -0.002155 | -0.002238 | -0.002354 | -0.002246 | -0.001962 | -0.001627 | -0.001384 | -0.001229 | -0.000992 | -0.000823 | -0.000712 | -0.000564 | -0.000444 | -0.00041 | -0.000462 | -0.000335 | -0.000289 | 0 | 0.000294 | 0.000574 | 0.000913 | 0.001329 | 0.001559 | 0.001668 | 0.001683 | 0.001585 | 0.001467 | 0.001292 | 0.001186 | 0.000938 | 0.000825
-0.002007 | -0.002144 | -0.002078 | -0.001934 | -0.001916 | -0.001953 | -0.002036 | -0.002126 | -0.002055 | -0.001785 | -0.00148 | -0.001268 | -0.001095 | -0.000872 | -0.000698 | -0.000618 | -0.000495 | -0.000374 | -0.000319 | -0.000424 | -0.000309 | -0.000264 | 0 | 0.00024 | 0.000434 | 0.000807 | 0.001194 | 0.001488 | 0.001584 | 0.001594 | 0.001467 | 0.001337 | 0.001152 | 0.001042 | 0.000821 | 0.000699
-0.001778 | -0.001874 | -0.001821 | -0.001664 | -0.001656 | -0.001707 | -0.00179 | -0.001902 | -0.001812 | -0.001576 | -0.001289 | -0.00108 | -0.00098 | -0.000772 | -0.000645 | -0.000637 | -0.000505 | -0.000387 | -0.00033 | -0.000401 | -0.000317 | -0.000238 | 0 | 0.000259 | 0.000504 | 0.000847 | 0.001232 | 0.001482 | 0.001531 | 0.001519 | 0.001402 | 0.001253 | 0.001094 | 0.00096 | 0.000724 | 0.000632
-0.001655 | -0.001719 | -0.001629 | -0.001526 | -0.001561 | -0.001641 | -0.001753 | -0.001913 | -0.001826 | -0.001633 | -0.001325 | -0.001129 | -0.000989 | -0.00078 | -0.000643 | -0.000602 | -0.000488 | -0.000368 | -0.000307 | -0.000395 | -0.000265 | -0.000267 | 0 | 0.000242 | 0.000452 | 0.000785 | 0.00118 | 0.001373 | 0.001374 | 0.001358 | 0.001201 | 0.001066 | 0.000895 | 0.000803 | 0.000542 | 0.000467
-0.001631 | -0.001703 | -0.001621 | -0.001487 | -0.001515 | -0.001586 | -0.001672 | -0.00184 | -0.001773 | -0.001554 | -0.001269 | -0.001061 | -0.000963 | -0.000751 | -0.000631 | -0.000594 | -0.000444 | -0.000325 | -0.000317 | -0.000392 | -0.000288 | -0.000219 | 0 | 0.000232 | 0.000438 | 0.000768 | 0.001143 | 0.001295 | 0.001276 | 0.001242 | 0.001067 | 0.00091 | 0.00076 | 0.00066 | 0.000435 | 0.000325
-0.001599 | -0.001613 | -0.001455 | -0.001323 | -0.001385 | -0.001447 | -0.001572 | -0.001684 | -0.001637 | -0.00142 | -0.001151 | -0.000947 | -0.000842 | -0.000655 | -0.000532 | -0.00049 | -0.00039 | -0.000301 | -0.000282 | -0.000334 | -0.000218 | -0.000189 | 0 | 0.000237 | 0.000418 | 0.000737 | 0.001061 | 0.001189 | 0.001148 | 0.001099 | 0.000948 | 0.000767 | 0.000614 | 0.000522 | 0.000288 | 0.000193
-0.001493 | -0.001491 | -0.001309 | -0.001185 | -0.001274 | -0.001392 | -0.001456 | -0.001648 | -0.001591 | -0.001386 | -0.001114 | -0.000901 | -0.000819 | -0.000627 | -0.000541 | -0.000456 | -0.000358 | -0.000255 | -0.000238 | -0.000331 | -0.000232 | -0.000191 | 0 | 0.000231 | 0.000393 | 0.000708 | 0.001018 | 0.001142 | 0.001054 | 0.000976 | 0.000809 | 0.000635 | 0.000515 | 0.000409 | 0.000155 | 6.5e-05
-0.001273 | -0.001227 | -0.001035 | -0.000945 | -0.00103 | -0.001137 | -0.001242 | -0.001437 | -0.001412 | -0.001245 | -0.001001 | -0.000824 | -0.000719 | -0.000536 | -0.00044 | -0.000411 | -0.000288 | -0.000265 | -0.000229 | -0.000322 | -0.000241 | -0.000182 | 0 | 0.000196 | 0.000349 | 0.000635 | 0.000917 | 0.000999 | 0.000892 | 0.000808 | 0.000643 | 0.000535 | 0.000333 | 0.000226 | 4e-06 | -7.5e-05
-0.001019 | -0.000972 | -0.000836 | -0.000776 | -0.000906 | -0.001054 | -0.00116 | -0.001372 | -0.001355 | -0.001186 | -0.000977 | -0.000786 | -0.000689 | -0.000524 | -0.000424 | -0.000368 | -0.000292 | -0.000227 | -0.000215 | -0.000296 | -0.00025 | -0.000161 | 0 | 0.000149 | 0.000314 | 0.000593 | 0.000863 | 0.000913 | 0.000815 | 0.00071 | 0.000543 | 0.000402 | 0.000246 | 0.000167 | -5.3e-05 | -0.000159
-0.000717 | -0.00067 | -0.000547 | -0.000534 | -0.000658 | -0.000843 | -0.000997 | -0.001181 | -0.001201 | -0.001058 | -0.000845 | -0.000689 | -0.000631 | -0.000466 | -0.00035 | -0.0003 | -0.000238 | -0.000166 | -0.000193 | -0.000245 | -0.000198 | -0.000143 | 0 | 0.000158 | 0.000301 | 0.000543 | 0.000832 | 0.000899 | 0.000797 | 0.000703 | 0.000507 | 0.000366 | 0.000189 | 0.000113 | -8.2e-05 | -0.00017
-1.8e-05 | 4e-05 | 8e-05 | 2.6e-05 | -0.000189 | -0.000381 | -0.000559 | -0.000779 | -0.000846 | -0.000746 | -0.00061 | -0.000455 | -0.000414 | -0.000269 | -0.000203 | -0.000177 | -0.000154 | -8.2e-05 | -7.3e-05 | -0.000177 | -0.000171 | -0.000123 | 0 | 0.000132 | 0.000245 | 0.00047 | 0.00072 | 0.000743 | 0.000605 | 0.00049 | 0.000355 | 0.000189 | 5.2e-05 | -1.5e-05 | -0.000224 | -0.000311
0.000328 | 0.00037 | 0.000381 | 0.000299 | 7.6e-05 | -0.00013 | -0.000361 | -0.000583 | -0.000639 | -0.000573 | -0.000437 | -0.000311 | -0.00028 | -0.000175 | -0.000125 | -0.000101 | -0.000129 | -2e-05 | -4.9e-05 | -0.000168 | -0.000146 | -0.000131 | 0 | 9.9e-05 | 0.000182 | 0.000382 | 0.000566 | 0.000592 | 0.000433 | 0.000347 | 0.000158 | 6e-06 | -0.000156 | -0.000208 | -0.00044 | -0.000503
0.000583 | 0.000603 | 0.000616 | 0.000493 | 0.000271 | 5.6e-05 | -0.000146 | -0.000366 | -0.00044 | -0.000382 | -0.000292 | -0.00018 | -0.000148 | -2.4e-05 | 2.3e-05 | 5e-06 | 1e-06 | 6.7e-05 | 6e-06 | -0.000146 | -9.8e-05 | -0.000132 | 0 | 8.4e-05 | 0.000128 | 0.000298 | 0.000447 | 0.000468 | 0.000275 | 0.000176 | -2e-05 | -0.00018 | -0.000299 | -0.000375 | -0.000591 | -0.000671
0.00085 | 0.000905 | 0.00092 | 0.0008 | 0.000551 | 0.000295 | 0.0001 | -0.000184 | -0.000233 | -0.000203 | -9.3e-05 | -2.9e-05 | 2.7e-05 | 7.1e-05 | 0.000105 | 8.1e-05 | 0.000121 | 8.1e-05 | 6.8e-05 | -0.000134 | -1.7e-05 | -3.3e-05 | 0 | 3.8e-05 | 4.1e-05 | 0.000173 | 0.000294 | 0.000252 | 6e-05 | -8.8e-05 | -0.000257 | -0.000394 | -0.000565 | -0.000614 | -0.00083 | -0.000905
0.001165 | 0.001219 | 0.001217 | 0.001092 | 0.000835 | 0.000596 | 0.000398 | 0.000158 | 7e-05 | 0.0001 | 0.00014 | 0.000186 | 0.000203 | 0.000281 | 0.000269 | 0.000236 | 0.000197 | 0.000195 | 6.9e-05 | -7.4e-05 | -5e-05 | -0.000112 | 0 | -1.4e-05 | -7.4e-05 | -2e-06 | 0.000104 | 2.3e-05 | -0.0002 | -0.000397 | -0.000556 | -0.000738 | -0.000858 | -0.000895 | -0.001166 | -0.001168
0.001469 | 0.001522 | 0.001522 | 0.001357 | 0.001074 | 0.000797 | 0.000648 | 0.000381 | 0.000283 | 0.00026 | 0.00032 | 0.000334 | 0.00037 | 0.000386 | 0.000375 | 0.000287 | 0.000278 | 0.000275 | 0.000155 | -4e-05 | -3.4e-05 | -5.8e-05 | 0 | -7.5e-05 | -0.000157 | -7.1e-05 | -6e-05 | -0.000156 | -0.000411 | -0.000538 | -0.000751 | -0.000902 | -0.001063 | -0.001081 | -0.001318 | -0.001376
0.001995 | 0.001994 | 0.001951 | 0.001776 | 0.001456 | 0.001218 | 0.001003 | 0.000726 | 0.000621 | 0.000601 | 0.000599 | 0.000591 | 0.000569 | 0.000611 | 0.000543 | 0.000451 | 0.000396 | 0.000368 | 0.000244 | -3.7e-05 | 5.8e-05 | -6.6e-05 | 0 | -9.8e-05 | -0.000248 | -0.000221 | -0.000244 | -0.000355 | -0.000635 | -0.000772 | -0.000998 | -0.001113 | -0.001249 | -0.001338 | -0.001564 | -0.001586
0.002374 | 0.00231 | 0.002228 | 0.002041 | 0.00173 | 0.001426 | 0.0013 | 0.001003 | 0.00084 | 0.000822 | 0.000777 | 0.000758 | 0.000727 | 0.000746 | 0.000665 | 0.000551 | 0.000518 | 0.000486 | 0.000352 | 5.1e-05 | 5.1e-05 | -6e-06 | 0 | -0.000126 | -0.000289 | -0.000267 | -0.000304 | -0.00042 | -0.00068 | -0.000863 | -0.001035 | -0.001175 | -0.001314 | -0.00135 | -0.001618 | -0.001663
0.002376 | 0.002303 | 0.002208 | 0.002059 | 0.001732 | 0.001486 | 0.001344 | 0.001078 | 0.0009 | 0.000908 | 0.000852 | 0.000846 | 0.000764 | 0.000818 | 0.000741 | 0.000602 | 0.000542 | 0.000507 | 0.000343 | 5.6e-05 | 8.1e-05 | 0 | 0 | -0.00014 | -0.000315 | -0.000337 | -0.000393 | -0.000505 | -0.000784 | -0.000952 | -0.001104 | -0.001267 | -0.001433 | -0.001431 | -0.001683 | -0.001732
0.002413 | 0.002326 | 0.002216 | 0.002039 | 0.001784 | 0.001524 | 0.001412 | 0.001107 | 0.00099 | 0.00093 | 0.000941 | 0.00089 | 0.000824 | 0.000831 | 0.000787 | 0.000662 | 0.000532 | 0.000567 | 0.000371 | 4.2e-05 | 0.000135 | 7e-06 | 0 | -0.000159 | -0.000314 | -0.00039 | -0.00043 | -0.00049 | -0.000803 | -0.000964 | -0.001071 | -0.001259 | -0.001393 | -0.001423 | -0.001647 | -0.001681
0.002166 | 0.00206 | 0.001992 | 0.001831 | 0.00161 | 0.001414 | 0.001265 | 0.001058 | 0.000952 | 0.000929 | 0.000922 | 0.000861 | 0.000804 | 0.000851 | 0.000761 | 0.000668 | 0.000558 | 0.000563 | 0.000331 | 5.7e-05 | 0.000111 | -2.9e-05 | 0 | -0.000137 | -0.000374 | -0.000404 | -0.00049 | -0.000503 | -0.00081 | -0.000957 | -0.001141 | -0.001207 | -0.001397 | -0.001378 | -0.001697 | -0.001699
0.001953 | 0.001831 | 0.001765 | 0.001663 | 0.001453 | 0.001262 | 0.00116 | 0.000971 | 0.000825 | 0.000849 | 0.000846 | 0.000798 | 0.000732 | 0.000769 | 0.000691 | 0.000606 | 0.000593 | 0.000572 | 0.000341 | -2e-06 | 6.6e-05 | -5.3e-05 | 0 | -0.000176 | -0.000371 | -0.000399 | -0.000459 | -0.000519 | -0.000818 | -0.000924 | -0.001098 | -0.001195 | -0.001345 | -0.001316 | -0.001617 | -0.001627
0.001754 | 0.001598 | 0.00152 | 0.001453 | 0.001248 | 0.001099 | 0.001029 | 0.00085 | 0.000705 | 0.000727 | 0.000763 | 0.000691 | 0.000691 | 0.000741 | 0.000668 | 0.000595 | 0.000547 | 0.000502 | 0.000334 | 2.3e-05 | 8e-06 | -3.7e-05 | 0 | -0.000183 | -0.000421 | -0.000408 | -0.000468 | -0.000528 | -0.000772 | -0.000934 | -0.001023 | -0.001102 | -0.001339 | -0.001312 | -0.001562 | -0.001565
0.001653 | 0.001481 | 0.001395 | 0.001295 | 0.001097 | 0.001012 | 0.000936 | 0.000727 | 0.000653 | 0.00068 | 0.000724 | 0.000683 | 0.000628 | 0.000704 | 0.000695 | 0.000551 | 0.000575 | 0.000524 | 0.000367 | -3e-06 | 6.6e-05 | -4.3e-05 | 0 | -0.000157 | -0.000329 | -0.000347 | -0.000368 | -0.000427 | -0.000644 | -0.000772 | -0.000869 | -0.000994 | -0.00116 | -0.001151 | -0.001445 | -0.001478
0.001434 | 0.001239 | 0.001151 | 0.001046 | 0.000899 | 0.000747 | 0.000759 | 0.000533 | 0.000469 | 0.000509 | 0.000578 | 0.000516 | 0.000515 | 0.000571 | 0.000618 | 0.000494 | 0.000503 | 0.000461 | 0.000256 | -7.1e-05 | 5.2e-05 | -6.9e-05 | 0 | -0.00012 | -0.000273 | -0.000272 | -0.000318 | -0.000291 | -0.000546 | -0.000667 | -0.000726 | -0.000867 | -0.001002 | -0.000969 | -0.001285 | -0.001253
0.001553 | 0.001293 | 0.001136 | 0.000978 | 0.00081 | 0.000651 | 0.000624 | 0.000401 | 0.000375 | 0.000418 | 0.000477 | 0.000456 | 0.000439 | 0.000559 | 0.000539 | 0.000389 | 0.000426 | 0.000492 | 0.000253 | -0.000163 | 1e-05 | -0.000127 | 0 | -0.000155 | -0.00026 | -0.000214 | -0.000218 | -0.000215 | -0.000371 | -0.000492 | -0.000582 | -0.000611 | -0.000845 | -0.00082 | -0.001068 | -0.001103
0.001704 | 0.001333 | 0.001104 | 0.000956 | 0.000698 | 0.000555 | 0.000534 | 0.000337 | 0.000241 | 0.0003 | 0.00042 | 0.000327 | 0.000293 | 0.000466 | 0.000547 | 0.000353 | 0.00045 | 0.000433 | 0.000249 | -0.000134 | 1.4e-05 | -8.4e-05 | 0 | -1.9e-05 | -0.000163 | -0.000103 | -6.7e-05 | 5e-06 | -0.000176 | -0.000278 | -0.000357 | -0.000387 | -0.000589 | -0.000505 | -0.000877 | -0.000892
0.001725 | 0.001305 | 0.000977 | 0.000761 | 0.000508 | 0.000365 | 0.000289 | 0.000108 | 8.9e-05 | 6e-05 | 0.000254 | 0.000149 | 8e-05 | 0.000294 | 0.000369 | 0.000185 | 0.000321 | 0.000326 | 0.000211 | -0.000204 | -3.3e-05 | -0.000244 | 0 | -4.5e-05 | -0.000108 | 2e-05 | 8e-06 | 0.000164 | -5e-05 | -7.3e-05 | -0.000109 | -0.00023 | -0.000399 | -0.000325 | -0.000659 | -0.000621

## ACS137_CC_tcarray
0.11464 | 0.10445 | 0.096266 | 0.089263 | 0.083497 | 0.078106 | 0.071799 | 0.066141 | 0.062409 | 0.057612 | 0.053196 | 0.047147 | 0.042855 | 0.037524 | 0.032814 | 0.026519 | 0.023535 | 0.018763 | 0.015219 | 0.010152 | 0.006575 | 0.004157 | 0 | -0.006807 | -0.013292 | -0.018555 | -0.023469 | -0.02877 | -0.034449 | -0.039154 | -0.043537 | -0.047454 | -0.049568 | -0.050928 | -0.052464 | -0.050856
0.11144 | 0.10046 | 0.092869 | 0.085464 | 0.079778 | 0.073704 | 0.067436 | 0.0628 | 0.059077 | 0.054591 | 0.050544 | 0.045735 | 0.040714 | 0.035363 | 0.030239 | 0.025279 | 0.022666 | 0.018357 | 0.014546 | 0.009764 | 0.006859 | 0.004823 | 0 | -0.005997 | -0.012352 | -0.01804 | -0.023263 | -0.028616 | -0.033631 | -0.03728 | -0.042643 | -0.044819 | -0.047823 | -0.048894 | -0.049348 | -0.048311
0.10566 | 0.095823 | 0.08783 | 0.08053 | 0.075622 | 0.070261 | 0.064484 | 0.059498 | 0.056589 | 0.05223 | 0.047989 | 0.04394 | 0.038384 | 0.034214 | 0.028952 | 0.02562 | 0.021429 | 0.016642 | 0.0136 | 0.009115 | 0.005824 | 0.003863 | 0 | -0.006716 | -0.012449 | -0.018373 | -0.022389 | -0.027467 | -0.03184 | -0.03608 | -0.040767 | -0.044896 | -0.04658 | -0.0474 | -0.048912 | -0.047612
0.10332 | 0.093796 | 0.084916 | 0.07838 | 0.07266 | 0.067863 | 0.063262 | 0.05855 | 0.054181 | 0.050179 | 0.045477 | 0.041213 | 0.036373 | 0.031968 | 0.027467 | 0.023611 | 0.019999 | 0.016262 | 0.013861 | 0.009037 | 0.005995 | 0.005219 | 0 | -0.005744 | -0.011848 | -0.016809 | -0.021905 | -0.02702 | -0.031143 | -0.035511 | -0.039466 | -0.041221 | -0.044298 | -0.045259 | -0.045403 | -0.044401
0.10008 | 0.091515 | 0.080588 | 0.07369 | 0.070341 | 0.065557 | 0.0588 | 0.055808 | 0.053992 | 0.049699 | 0.045676 | 0.040613 | 0.036652 | 0.032754 | 0.02795 | 0.023972 | 0.01999 | 0.016065 | 0.013049 | 0.01006 | 0.005823 | 0.003403 | 0 | -0.00558 | -0.01158 | -0.017116 | -0.021016 | -0.025884 | -0.029789 | -0.033547 | -0.038285 | -0.041579 | -0.042953 | -0.044101 | -0.045 | -0.043918
0.097518 | 0.087351 | 0.079364 | 0.072777 | 0.067177 | 0.062549 | 0.058268 | 0.054385 | 0.050685 | 0.046723 | 0.042751 | 0.038802 | 0.033691 | 0.029845 | 0.025507 | 0.021875 | 0.01831 | 0.015342 | 0.012319 | 0.00841 | 0.005717 | 0.004134 | 0 | -0.005568 | -0.01133 | -0.016575 | -0.020758 | -0.025185 | -0.029535 | -0.0328 | -0.036585 | -0.038934 | -0.040779 | -0.041502 | -0.041984 | -0.040933
0.094977 | 0.085564 | 0.07675 | 0.070474 | 0.065606 | 0.061129 | 0.056669 | 0.05282 | 0.049665 | 0.04517 | 0.041516 | 0.037438 | 0.033292 | 0.029096 | 0.024918 | 0.021709 | 0.018543 | 0.01458 | 0.012203 | 0.008553 | 0.005871 | 0.004106 | 0 | -0.005154 | -0.010765 | -0.015929 | -0.020134 | -0.024374 | -0.028497 | -0.031731 | -0.035077 | -0.037424 | -0.039337 | -0.040091 | -0.040479 | -0.039729
0.092086 | 0.082905 | 0.074322 | 0.068234 | 0.063263 | 0.058761 | 0.054764 | 0.051077 | 0.047492 | 0.043557 | 0.03972 | 0.035686 | 0.031494 | 0.027645 | 0.02362 | 0.020647 | 0.017521 | 0.01422 | 0.01145 | 0.008161 | 0.005129 | 0.003651 | 0 | -0.005163 | -0.010846 | -0.015398 | -0.019863 | -0.023898 | -0.027562 | -0.03038 | -0.03343 | -0.03574 | -0.037681 | -0.038257 | -0.039101 | -0.038158
0.088232 | 0.079155 | 0.071256 | 0.064994 | 0.060574 | 0.05654 | 0.052648 | 0.048962 | 0.045702 | 0.042301 | 0.038384 | 0.034287 | 0.030828 | 0.026938 | 0.02316 | 0.019856 | 0.016974 | 0.013747 | 0.01093 | 0.007802 | 0.004804 | 0.003594 | 0 | -0.004932 | -0.010239 | -0.015289 | -0.019036 | -0.02282 | -0.026025 | -0.028802 | -0.031647 | -0.033698 | -0.035534 | -0.03621 | -0.036461 | -0.036016
0.084955 | 0.076467 | 0.069276 | 0.063294 | 0.058804 | 0.054769 | 0.050772 | 0.047516 | 0.044311 | 0.040726 | 0.037389 | 0.033218 | 0.029761 | 0.02595 | 0.022386 | 0.019367 | 0.016471 | 0.01332 | 0.010743 | 0.00749 | 0.005235 | 0.003537 | 0 | -0.004487 | -0.009959 | -0.014281 | -0.017841 | -0.021376 | -0.024422 | -0.026984 | -0.029615 | -0.031701 | -0.033358 | -0.034037 | -0.034649 | -0.033876
0.082607 | 0.073966 | 0.067008 | 0.061271 | 0.057228 | 0.053231 | 0.049122 | 0.045919 | 0.043064 | 0.03976 | 0.036219 | 0.032158 | 0.028912 | 0.025388 | 0.021942 | 0.018498 | 0.016019 | 0.012847 | 0.010279 | 0.00747 | 0.004893 | 0.003027 | 0 | -0.004727 | -0.00955 | -0.013913 | -0.01747 | -0.020497 | -0.022845 | -0.02501 | -0.02791 | -0.029927 | -0.031545 | -0.032053 | -0.032985 | -0.03226
0.079914 | 0.071437 | 0.064777 | 0.059085 | 0.055191 | 0.05129 | 0.047472 | 0.044487 | 0.041542 | 0.037995 | 0.03498 | 0.031236 | 0.027771 | 0.024504 | 0.021127 | 0.017953 | 0.015506 | 0.012504 | 0.01001 | 0.006884 | 0.004613 | 0.00337 | 0 | -0.004452 | -0.009509 | -0.013314 | -0.016552 | -0.019534 | -0.02171 | -0.024136 | -0.026223 | -0.028113 | -0.029489 | -0.030222 | -0.030807 | -0.030206
0.078973 | 0.070452 | 0.064145 | 0.058583 | 0.054524 | 0.050715 | 0.047187 | 0.043957 | 0.040802 | 0.037523 | 0.033923 | 0.030212 | 0.027007 | 0.023673 | 0.020354 | 0.017664 | 0.014973 | 0.012189 | 0.009569 | 0.006863 | 0.004467 | 0.003154 | 0 | -0.00452 | -0.009099 | -0.013057 | -0.015979 | -0.018533 | -0.020834 | -0.022811 | -0.025076 | -0.026713 | -0.027903 | -0.028601 | -0.028943 | -0.028332
0.076025 | 0.067726 | 0.061524 | 0.056286 | 0.05224 | 0.048899 | 0.045227 | 0.04218 | 0.039247 | 0.036162 | 0.03282 | 0.029053 | 0.026025 | 0.023045 | 0.019788 | 0.01712 | 0.014539 | 0.011549 | 0.009182 | 0.006652 | 0.004321 | 0.002845 | 0 | -0.004408 | -0.008592 | -0.012349 | -0.015368 | -0.017708 | -0.019503 | -0.021318 | -0.023456 | -0.025095 | -0.026148 | -0.026601 | -0.027242 | -0.026426
0.074903 | 0.066584 | 0.060489 | 0.055169 | 0.050986 | 0.047386 | 0.043853 | 0.041024 | 0.037971 | 0.034769 | 0.031343 | 0.028022 | 0.024845 | 0.021765 | 0.018922 | 0.016128 | 0.013828 | 0.011115 | 0.008889 | 0.006358 | 0.003909 | 0.00281 | 0 | -0.00395 | -0.008229 | -0.011784 | -0.014611 | -0.016797 | -0.01861 | -0.020275 | -0.022143 | -0.023591 | -0.024732 | -0.025142 | -0.02529 | -0.024596
0.073432 | 0.065277 | 0.059227 | 0.053816 | 0.04976 | 0.046245 | 0.042933 | 0.039774 | 0.036639 | 0.033475 | 0.030218 | 0.026811 | 0.023875 | 0.020985 | 0.017929 | 0.015531 | 0.013237 | 0.010764 | 0.008449 | 0.00579 | 0.003742 | 0.002703 | 0 | -0.003831 | -0.007841 | -0.011238 | -0.013952 | -0.015936 | -0.017387 | -0.019097 | -0.020927 | -0.022167 | -0.023013 | -0.023403 | -0.023505 | -0.022771
0.072954 | 0.064512 | 0.058541 | 0.053217 | 0.049075 | 0.045513 | 0.042202 | 0.038954 | 0.035893 | 0.03282 | 0.029444 | 0.026055 | 0.023223 | 0.020343 | 0.017459 | 0.015042 | 0.013017 | 0.010391 | 0.008415 | 0.005938 | 0.003517 | 0.002493 | 0 | -0.003662 | -0.007723 | -0.011056 | -0.013617 | -0.015554 | -0.016893 | -0.018307 | -0.019707 | -0.021308 | -0.021964 | -0.022231 | -0.022203 | -0.021378
0.072393 | 0.06412 | 0.057986 | 0.052478 | 0.048466 | 0.044962 | 0.041243 | 0.03843 | 0.035437 | 0.032167 | 0.028842 | 0.025586 | 0.022822 | 0.020014 | 0.01714 | 0.014787 | 0.012506 | 0.010336 | 0.007968 | 0.00565 | 0.003387 | 0.002515 | 0 | -0.003528 | -0.007379 | -0.010741 | -0.013178 | -0.015001 | -0.016172 | -0.017582 | -0.019231 | -0.020217 | -0.020935 | -0.021224 | -0.021115 | -0.020163
0.071643 | 0.063247 | 0.057118 | 0.051681 | 0.04758 | 0.044089 | 0.040562 | 0.037569 | 0.034522 | 0.031216 | 0.028124 | 0.024805 | 0.022156 | 0.019422 | 0.016768 | 0.0143 | 0.012206 | 0.009978 | 0.007919 | 0.005516 | 0.003238 | 0.002441 | 0 | -0.003483 | -0.007196 | -0.010234 | -0.012668 | -0.014215 | -0.015421 | -0.016761 | -0.018061 | -0.019249 | -0.01981 | -0.020029 | -0.019913 | -0.019017
0.070187 | 0.062026 | 0.055957 | 0.050631 | 0.046597 | 0.043141 | 0.039797 | 0.036838 | 0.033908 | 0.030679 | 0.027414 | 0.024517 | 0.021695 | 0.019103 | 0.01649 | 0.014115 | 0.012098 | 0.009854 | 0.007741 | 0.00537 | 0.003307 | 0.002367 | 0 | -0.003296 | -0.006998 | -0.010157 | -0.012247 | -0.01381 | -0.014786 | -0.016128 | -0.017417 | -0.018465 | -0.019064 | -0.019107 | -0.019126 | -0.018008
0.069302 | 0.06106 | 0.055031 | 0.049613 | 0.045523 | 0.042089 | 0.038525 | 0.03573 | 0.032785 | 0.029645 | 0.02664 | 0.023494 | 0.020972 | 0.018434 | 0.015862 | 0.013657 | 0.011676 | 0.009417 | 0.007385 | 0.00511 | 0.002905 | 0.002234 | 0 | -0.00325 | -0.006763 | -0.009751 | -0.01182 | -0.01332 | -0.014109 | -0.015469 | -0.016704 | -0.017768 | -0.018253 | -0.018259 | -0.018053 | -0.017052
0.068472 | 0.060447 | 0.054481 | 0.049122 | 0.045226 | 0.041865 | 0.038251 | 0.03549 | 0.032567 | 0.029415 | 0.026345 | 0.023323 | 0.020759 | 0.018279 | 0.015757 | 0.013501 | 0.011482 | 0.009311 | 0.007405 | 0.005166 | 0.003055 | 0.002408 | 0 | -0.003041 | -0.006472 | -0.009501 | -0.011523 | -0.012851 | -0.013663 | -0.014868 | -0.016043 | -0.017023 | -0.017544 | -0.017576 | -0.017315 | -0.016406
0.067326 | 0.059398 | 0.053581 | 0.048334 | 0.044278 | 0.040912 | 0.037695 | 0.034752 | 0.031846 | 0.028913 | 0.025833 | 0.022837 | 0.020275 | 0.017867 | 0.015443 | 0.013146 | 0.011367 | 0.009074 | 0.007281 | 0.004839 | 0.002897 | 0.002141 | 0 | -0.003052 | -0.006431 | -0.009303 | -0.01116 | -0.012542 | -0.013313 | -0.014569 | -0.015645 | -0.016494 | -0.016995 | -0.016916 | -0.01665 | -0.015685
0.06624 | 0.058498 | 0.052728 | 0.047595 | 0.04372 | 0.040384 | 0.037067 | 0.034343 | 0.031533 | 0.028629 | 0.025564 | 0.022711 | 0.020147 | 0.017746 | 0.015264 | 0.013108 | 0.011079 | 0.009038 | 0.00712 | 0.004889 | 0.002834 | 0.002231 | 0 | -0.002995 | -0.006376 | -0.009202 | -0.011169 | -0.012428 | -0.013159 | -0.014476 | -0.015512 | -0.016513 | -0.016865 | -0.016766 | -0.016571 | -0.015491
0.065568 | 0.057998 | 0.052323 | 0.04723 | 0.043509 | 0.040316 | 0.03693 | 0.03432 | 0.031568 | 0.028576 | 0.025626 | 0.022719 | 0.020209 | 0.0177 | 0.015207 | 0.013092 | 0.011114 | 0.009111 | 0.00712 | 0.004953 | 0.002842 | 0.002302 | 0 | -0.002909 | -0.006255 | -0.009098 | -0.011009 | -0.012259 | -0.012969 | -0.014269 | -0.015276 | -0.01627 | -0.016715 | -0.01658 | -0.016399 | -0.015267
0.064459 | 0.057006 | 0.05145 | 0.046492 | 0.042781 | 0.039686 | 0.036498 | 0.033844 | 0.031109 | 0.028306 | 0.025346 | 0.022551 | 0.019964 | 0.017567 | 0.015138 | 0.012997 | 0.011032 | 0.008924 | 0.006998 | 0.004895 | 0.002775 | 0.00223 | 0 | -0.002993 | -0.006393 | -0.009186 | -0.011141 | -0.01245 | -0.013287 | -0.014601 | -0.015688 | -0.01666 | -0.016981 | -0.016866 | -0.016551 | -0.015446
0.063698 | 0.056315 | 0.050891 | 0.046023 | 0.04226 | 0.039178 | 0.036126 | 0.033537 | 0.03081 | 0.027952 | 0.025101 | 0.022397 | 0.019814 | 0.017405 | 0.014971 | 0.012853 | 0.010921 | 0.008875 | 0.007063 | 0.004807 | 0.002661 | 0.002163 | 0 | -0.003051 | -0.006384 | -0.009243 | -0.011155 | -0.012602 | -0.013494 | -0.014876 | -0.015923 | -0.016892 | -0.017251 | -0.017145 | -0.016881 | -0.015762
0.062205 | 0.055154 | 0.049976 | 0.045304 | 0.041803 | 0.038815 | 0.035689 | 0.033205 | 0.030606 | 0.027792 | 0.024974 | 0.022247 | 0.019695 | 0.017376 | 0.014979 | 0.012832 | 0.010847 | 0.008854 | 0.007003 | 0.004871 | 0.002744 | 0.002322 | 0 | -0.002913 | -0.006234 | -0.009113 | -0.011065 | -0.012634 | -0.013559 | -0.014993 | -0.016162 | -0.01714 | -0.017519 | -0.017388 | -0.017071 | -0.015906
0.061772 | 0.054821 | 0.04966 | 0.045021 | 0.041483 | 0.038488 | 0.035445 | 0.032987 | 0.030423 | 0.027675 | 0.024896 | 0.022216 | 0.019669 | 0.017229 | 0.014858 | 0.012757 | 0.010897 | 0.008801 | 0.006989 | 0.004796 | 0.002694 | 0.002299 | 0 | -0.002936 | -0.006269 | -0.009128 | -0.011135 | -0.012789 | -0.013928 | -0.015419 | -0.016611 | -0.017627 | -0.017918 | -0.017823 | -0.017489 | -0.016376
0.060983 | 0.054232 | 0.049178 | 0.044658 | 0.041205 | 0.038286 | 0.035284 | 0.032756 | 0.030145 | 0.027468 | 0.024721 | 0.022015 | 0.019517 | 0.017077 | 0.014669 | 0.012565 | 0.010648 | 0.008716 | 0.007011 | 0.004761 | 0.002705 | 0.002193 | 0 | -0.002934 | -0.006258 | -0.00908 | -0.011123 | -0.012861 | -0.014151 | -0.015763 | -0.016979 | -0.017984 | -0.018348 | -0.018147 | -0.017792 | -0.016698
0.059484 | 0.053035 | 0.048197 | 0.043814 | 0.040463 | 0.037607 | 0.034721 | 0.032311 | 0.029866 | 0.027262 | 0.024523 | 0.021849 | 0.019431 | 0.017047 | 0.014611 | 0.012535 | 0.010686 | 0.00867 | 0.006854 | 0.004747 | 0.002673 | 0.002279 | 0 | -0.002858 | -0.006218 | -0.008987 | -0.011128 | -0.013011 | -0.014296 | -0.016053 | -0.017281 | -0.018286 | -0.018668 | -0.018533 | -0.018204 | -0.017065
0.05885 | 0.052488 | 0.047754 | 0.043427 | 0.040109 | 0.037374 | 0.034445 | 0.032077 | 0.02966 | 0.027053 | 0.024398 | 0.021709 | 0.0193 | 0.016932 | 0.014536 | 0.012413 | 0.010566 | 0.008521 | 0.006787 | 0.004643 | 0.002667 | 0.002113 | 0 | -0.00293 | -0.006189 | -0.008939 | -0.011172 | -0.013111 | -0.014602 | -0.016351 | -0.017751 | -0.018858 | -0.01914 | -0.019051 | -0.018646 | -0.017538
0.058035 | 0.051857 | 0.047178 | 0.042895 | 0.039647 | 0.03685 | 0.03401 | 0.031666 | 0.029301 | 0.02674 | 0.024136 | 0.021548 | 0.019136 | 0.016817 | 0.014388 | 0.012347 | 0.01052 | 0.008493 | 0.006845 | 0.004687 | 0.002678 | 0.002129 | 0 | -0.002839 | -0.006094 | -0.00885 | -0.011071 | -0.013184 | -0.014773 | -0.016675 | -0.018122 | -0.019162 | -0.019561 | -0.019342 | -0.019029 | -0.017956
0.056728 | 0.050643 | 0.046113 | 0.04197 | 0.038732 | 0.036043 | 0.033316 | 0.031072 | 0.028713 | 0.026205 | 0.023635 | 0.021094 | 0.018751 | 0.016516 | 0.014167 | 0.012095 | 0.010305 | 0.008397 | 0.006571 | 0.004526 | 0.00252 | 0.002059 | 0 | -0.002834 | -0.006081 | -0.00889 | -0.011112 | -0.013333 | -0.014961 | -0.016937 | -0.018367 | -0.019504 | -0.019788 | -0.019692 | -0.019347 | -0.018258
0.056006 | 0.049946 | 0.045439 | 0.041299 | 0.038098 | 0.035451 | 0.032768 | 0.030525 | 0.028236 | 0.025787 | 0.02324 | 0.020821 | 0.01847 | 0.016273 | 0.013962 | 0.011984 | 0.010218 | 0.008282 | 0.006572 | 0.004514 | 0.002539 | 0.002132 | 0 | -0.002778 | -0.005972 | -0.008758 | -0.011014 | -0.013234 | -0.015051 | -0.017099 | -0.018587 | -0.019673 | -0.020032 | -0.019909 | -0.019641 | -0.018557
0.054544 | 0.048757 | 0.044363 | 0.040319 | 0.037211 | 0.034619 | 0.031995 | 0.029794 | 0.027591 | 0.025238 | 0.02279 | 0.020359 | 0.018133 | 0.015935 | 0.013652 | 0.011747 | 0.009959 | 0.008097 | 0.00642 | 0.00445 | 0.002429 | 0.002029 | 0 | -0.002735 | -0.005869 | -0.008657 | -0.010854 | -0.013103 | -0.014865 | -0.016909 | -0.018438 | -0.019549 | -0.019943 | -0.019827 | -0.019485 | -0.018413
0.053428 | 0.047626 | 0.043293 | 0.039322 | 0.036239 | 0.033668 | 0.031106 | 0.028951 | 0.026744 | 0.024463 | 0.022031 | 0.019755 | 0.017539 | 0.015504 | 0.013234 | 0.011405 | 0.009742 | 0.007945 | 0.006274 | 0.004264 | 0.002453 | 0.002048 | 0 | -0.002656 | -0.005697 | -0.008391 | -0.010554 | -0.012785 | -0.014612 | -0.016674 | -0.018188 | -0.019247 | -0.019589 | -0.019472 | -0.019141 | -0.018096
0.052255 | 0.046575 | 0.042282 | 0.038369 | 0.035349 | 0.032847 | 0.030326 | 0.028261 | 0.026137 | 0.023923 | 0.021582 | 0.019325 | 0.017208 | 0.015195 | 0.012986 | 0.011222 | 0.009509 | 0.007727 | 0.006067 | 0.004219 | 0.002272 | 0.001963 | 0 | -0.00255 | -0.005546 | -0.008209 | -0.010382 | -0.012499 | -0.014288 | -0.016335 | -0.017832 | -0.018921 | -0.019261 | -0.019132 | -0.018796 | -0.017731
0.051348 | 0.045695 | 0.041503 | 0.037647 | 0.034677 | 0.032217 | 0.029748 | 0.027661 | 0.025574 | 0.023418 | 0.021184 | 0.01898 | 0.016921 | 0.014864 | 0.012714 | 0.010989 | 0.009328 | 0.007546 | 0.006008 | 0.004056 | 0.00222 | 0.001887 | 0 | -0.002554 | -0.005504 | -0.00803 | -0.010169 | -0.012284 | -0.014022 | -0.016062 | -0.01756 | -0.018585 | -0.018898 | -0.018786 | -0.018398 | -0.017405
0.050941 | 0.045308 | 0.041086 | 0.037257 | 0.034263 | 0.031779 | 0.029336 | 0.02739 | 0.025289 | 0.023183 | 0.020922 | 0.018794 | 0.016734 | 0.014757 | 0.012644 | 0.010912 | 0.009256 | 0.007521 | 0.005927 | 0.004039 | 0.002179 | 0.001877 | 0 | -0.002472 | -0.005364 | -0.007918 | -0.010008 | -0.012045 | -0.013781 | -0.015833 | -0.017216 | -0.018237 | -0.018551 | -0.018402 | -0.018019 | -0.016965
0.050456 | 0.044792 | 0.040575 | 0.036757 | 0.033775 | 0.031333 | 0.028927 | 0.026928 | 0.024879 | 0.02277 | 0.020597 | 0.01847 | 0.016504 | 0.014525 | 0.012419 | 0.010724 | 0.009131 | 0.007403 | 0.00581 | 0.004002 | 0.002151 | 0.001857 | 0 | -0.002442 | -0.005321 | -0.007861 | -0.009938 | -0.011968 | -0.013614 | -0.015591 | -0.016999 | -0.018018 | -0.018285 | -0.018123 | -0.017725 | -0.016688
0.049521 | 0.043926 | 0.039773 | 0.035989 | 0.033118 | 0.030707 | 0.028291 | 0.026364 | 0.02437 | 0.022309 | 0.02019 | 0.018136 | 0.016191 | 0.014305 | 0.012255 | 0.0106 | 0.009011 | 0.007317 | 0.005791 | 0.003939 | 0.002116 | 0.001812 | 0 | -0.002413 | -0.00524 | -0.007774 | -0.009838 | -0.011835 | -0.013521 | -0.015433 | -0.016814 | -0.017843 | -0.018108 | -0.017921 | -0.017544 | -0.016479
0.052349 | 0.046598 | 0.042398 | 0.038832 | 0.03614 | 0.033941 | 0.031655 | 0.029826 | 0.027852 | 0.025798 | 0.02355 | 0.021283 | 0.019037 | 0.016819 | 0.014468 | 0.01244 | 0.010643 | 0.008659 | 0.006912 | 0.004794 | 0.002769 | 0.002227 | 0 | -0.003093 | -0.0065 | -0.009591 | -0.012191 | -0.014664 | -0.016624 | -0.018762 | -0.020242 | -0.021257 | -0.021448 | -0.021227 | -0.020732 | -0.019583
0.052712 | 0.047047 | 0.042888 | 0.039353 | 0.0367 | 0.034439 | 0.032135 | 0.030293 | 0.028272 | 0.026207 | 0.023946 | 0.021618 | 0.019353 | 0.017105 | 0.014714 | 0.01266 | 0.010821 | 0.00883 | 0.007014 | 0.004889 | 0.002834 | 0.002286 | 0 | -0.003154 | -0.006677 | -0.009838 | -0.012514 | -0.015063 | -0.017027 | -0.019191 | -0.02065 | -0.021686 | -0.021873 | -0.021637 | -0.021134 | -0.019982
0.052966 | 0.047397 | 0.043329 | 0.039844 | 0.037173 | 0.034934 | 0.032684 | 0.030781 | 0.02875 | 0.026678 | 0.024369 | 0.022021 | 0.019724 | 0.017404 | 0.014972 | 0.012897 | 0.010997 | 0.00896 | 0.007156 | 0.004985 | 0.002908 | 0.002297 | 0 | -0.003246 | -0.006821 | -0.010048 | -0.012787 | -0.015356 | -0.017364 | -0.019575 | -0.021029 | -0.022024 | -0.022212 | -0.021954 | -0.021416 | -0.020264
0.05373 | 0.048185 | 0.044082 | 0.040533 | 0.037811 | 0.035543 | 0.033199 | 0.031261 | 0.029211 | 0.027039 | 0.024649 | 0.022234 | 0.019903 | 0.017603 | 0.015134 | 0.013023 | 0.011108 | 0.009095 | 0.007218 | 0.005053 | 0.002965 | 0.002372 | 0 | -0.003275 | -0.006905 | -0.010142 | -0.01292 | -0.015539 | -0.017525 | -0.019728 | -0.0212 | -0.022192 | -0.022379 | -0.022112 | -0.021574 | -0.020386
0.054422 | 0.048927 | 0.044888 | 0.041298 | 0.038545 | 0.03621 | 0.033837 | 0.031872 | 0.029802 | 0.027623 | 0.025174 | 0.022717 | 0.020305 | 0.0179 | 0.015369 | 0.013183 | 0.011248 | 0.0092 | 0.007283 | 0.005113 | 0.003031 | 0.002403 | 0 | -0.003354 | -0.007043 | -0.010367 | -0.013206 | -0.015851 | -0.017913 | -0.020083 | -0.02158 | -0.022588 | -0.022745 | -0.022464 | -0.021917 | -0.020745
0.055028 | 0.049657 | 0.045627 | 0.041938 | 0.03903 | 0.036657 | 0.034233 | 0.032206 | 0.030056 | 0.027738 | 0.025242 | 0.022712 | 0.0203 | 0.017928 | 0.015377 | 0.013192 | 0.011265 | 0.009256 | 0.007322 | 0.005156 | 0.003051 | 0.002426 | 0 | -0.003314 | -0.007005 | -0.010279 | -0.013109 | -0.015752 | -0.017813 | -0.02003 | -0.021481 | -0.022434 | -0.022581 | -0.022315 | -0.021745 | -0.020561
0.055216 | 0.049785 | 0.045748 | 0.042037 | 0.039119 | 0.036685 | 0.034249 | 0.032255 | 0.030097 | 0.027765 | 0.025234 | 0.022691 | 0.020267 | 0.017831 | 0.015311 | 0.013125 | 0.011173 | 0.009186 | 0.007272 | 0.005102 | 0.002976 | 0.002405 | 0 | -0.003281 | -0.006951 | -0.01024 | -0.013054 | -0.015692 | -0.017689 | -0.019849 | -0.021307 | -0.022278 | -0.022401 | -0.022095 | -0.021534 | -0.02038
0.055146 | 0.049691 | 0.045603 | 0.041834 | 0.038852 | 0.036373 | 0.033904 | 0.031858 | 0.02971 | 0.027383 | 0.024828 | 0.022345 | 0.019927 | 0.017564 | 0.015046 | 0.01292 | 0.011003 | 0.009008 | 0.007189 | 0.005027 | 0.002953 | 0.002378 | 0 | -0.003236 | -0.006836 | -0.010076 | -0.012816 | -0.01543 | -0.017415 | -0.019498 | -0.020947 | -0.021907 | -0.021974 | -0.021684 | -0.021127 | -0.019935
0.054336 | 0.048856 | 0.044728 | 0.040992 | 0.03803 | 0.035615 | 0.033185 | 0.031173 | 0.029061 | 0.026754 | 0.024249 | 0.02182 | 0.019396 | 0.017059 | 0.014559 | 0.01251 | 0.010664 | 0.008707 | 0.006908 | 0.004823 | 0.0028 | 0.002242 | 0 | -0.003116 | -0.006579 | -0.009736 | -0.01244 | -0.014973 | -0.01689 | -0.018987 | -0.020398 | -0.021296 | -0.021397 | -0.021065 | -0.020491 | -0.019289
0.05336 | 0.047925 | 0.043845 | 0.040136 | 0.037259 | 0.03487 | 0.032481 | 0.030533 | 0.028459 | 0.026235 | 0.023735 | 0.021365 | 0.019059 | 0.016766 | 0.014318 | 0.012329 | 0.010499 | 0.008589 | 0.006831 | 0.004782 | 0.002751 | 0.002264 | 0 | -0.003039 | -0.006459 | -0.009478 | -0.012084 | -0.014505 | -0.016368 | -0.018371 | -0.019703 | -0.020585 | -0.020656 | -0.020308 | -0.019732 | -0.018555
0.052211 | 0.04667 | 0.042533 | 0.03899 | 0.036254 | 0.033976 | 0.031678 | 0.029801 | 0.027762 | 0.025629 | 0.023256 | 0.020949 | 0.018651 | 0.016446 | 0.014023 | 0.012081 | 0.010308 | 0.008423 | 0.006673 | 0.004667 | 0.002688 | 0.002184 | 0 | -0.002974 | -0.006292 | -0.009271 | -0.011833 | -0.014226 | -0.01606 | -0.01804 | -0.01936 | -0.020231 | -0.020313 | -0.019952 | -0.019348 | -0.018151
0.050759 | 0.045279 | 0.041143 | 0.037652 | 0.035056 | 0.032876 | 0.030639 | 0.028846 | 0.026882 | 0.024856 | 0.02257 | 0.020314 | 0.018159 | 0.01598 | 0.013649 | 0.011763 | 0.009991 | 0.008171 | 0.006478 | 0.004545 | 0.002553 | 0.00214 | 0 | -0.002904 | -0.006188 | -0.009119 | -0.011601 | -0.013951 | -0.01573 | -0.017727 | -0.018997 | -0.019874 | -0.01996 | -0.019597 | -0.019012 | -0.017796
0.049699 | 0.044172 | 0.040072 | 0.036653 | 0.034166 | 0.032091 | 0.029984 | 0.028214 | 0.0263 | 0.024297 | 0.022104 | 0.019916 | 0.017776 | 0.015661 | 0.01341 | 0.011569 | 0.009872 | 0.008095 | 0.006435 | 0.004472 | 0.002543 | 0.002108 | 0 | -0.00283 | -0.006051 | -0.008891 | -0.011328 | -0.013629 | -0.015355 | -0.017278 | -0.018555 | -0.019382 | -0.019429 | -0.019077 | -0.018473 | -0.017281
0.049028 | 0.043308 | 0.039228 | 0.035937 | 0.033499 | 0.031447 | 0.029396 | 0.027708 | 0.025869 | 0.023899 | 0.021712 | 0.019588 | 0.017496 | 0.015417 | 0.013182 | 0.011381 | 0.009688 | 0.007974 | 0.006345 | 0.004406 | 0.002528 | 0.002105 | 0 | -0.002805 | -0.005972 | -0.008801 | -0.011145 | -0.01343 | -0.015136 | -0.017047 | -0.018262 | -0.019089 | -0.019124 | -0.018772 | -0.018156 | -0.016967
0.048057 | 0.042606 | 0.038561 | 0.035332 | 0.032966 | 0.030972 | 0.028965 | 0.027302 | 0.025458 | 0.023571 | 0.02148 | 0.019366 | 0.017328 | 0.015272 | 0.013076 | 0.011302 | 0.009592 | 0.007874 | 0.006279 | 0.004353 | 0.002455 | 0.002064 | 0 | -0.002764 | -0.00588 | -0.008686 | -0.011051 | -0.013259 | -0.014915 | -0.016761 | -0.017961 | -0.01876 | -0.018769 | -0.018449 | -0.017842 | -0.016675
0.047394 | 0.041977 | 0.037913 | 0.034748 | 0.032412 | 0.030465 | 0.0285 | 0.026896 | 0.025091 | 0.023237 | 0.021192 | 0.0191 | 0.017075 | 0.015082 | 0.012874 | 0.011076 | 0.009448 | 0.00772 | 0.006138 | 0.004279 | 0.002406 | 0.002013 | 0 | -0.002746 | -0.005825 | -0.008574 | -0.010907 | -0.013055 | -0.014696 | -0.016512 | -0.017662 | -0.018441 | -0.018485 | -0.018113 | -0.017524 | -0.016372
0.046949 | 0.04147 | 0.037429 | 0.034306 | 0.032019 | 0.030109 | 0.028121 | 0.026529 | 0.024804 | 0.023001 | 0.020978 | 0.01893 | 0.016908 | 0.014904 | 0.012758 | 0.011033 | 0.00939 | 0.007728 | 0.006152 | 0.004246 | 0.002383 | 0.001987 | 0 | -0.002729 | -0.005835 | -0.008573 | -0.010874 | -0.013009 | -0.014579 | -0.016371 | -0.017482 | -0.018272 | -0.0183 | -0.017929 | -0.017319 | -0.016211
0.046358 | 0.040949 | 0.036949 | 0.033865 | 0.031673 | 0.02982 | 0.027939 | 0.026377 | 0.024688 | 0.022897 | 0.020891 | 0.018825 | 0.016872 | 0.014874 | 0.012731 | 0.010974 | 0.009387 | 0.007687 | 0.006082 | 0.004238 | 0.002415 | 0.001977 | 0 | -0.002725 | -0.005779 | -0.008509 | -0.010781 | -0.012847 | -0.014414 | -0.016143 | -0.017267 | -0.018048 | -0.018047 | -0.017685 | -0.017118 | -0.015982
0.046006 | 0.040664 | 0.036701 | 0.033668 | 0.031444 | 0.029621 | 0.027739 | 0.026205 | 0.024489 | 0.022738 | 0.020763 | 0.018733 | 0.016757 | 0.01478 | 0.012617 | 0.010909 | 0.009319 | 0.007606 | 0.006034 | 0.004186 | 0.002378 | 0.001959 | 0 | -0.002688 | -0.005724 | -0.008397 | -0.010656 | -0.012717 | -0.014228 | -0.015917 | -0.01701 | -0.017772 | -0.017799 | -0.017427 | -0.016834 | -0.015731
0.045657 | 0.04028 | 0.036318 | 0.033338 | 0.031167 | 0.02934 | 0.02747 | 0.025966 | 0.024295 | 0.022531 | 0.020571 | 0.018576 | 0.016614 | 0.014693 | 0.012556 | 0.010837 | 0.009276 | 0.007609 | 0.005996 | 0.00419 | 0.002404 | 0.001957 | 0 | -0.00267 | -0.005697 | -0.008363 | -0.01061 | -0.012615 | -0.014131 | -0.015808 | -0.016933 | -0.017638 | -0.017688 | -0.017311 | -0.016749 | -0.01565
0.044888 | 0.039633 | 0.035806 | 0.032889 | 0.030789 | 0.028991 | 0.027181 | 0.025728 | 0.02406 | 0.022333 | 0.020376 | 0.018406 | 0.016462 | 0.014553 | 0.012446 | 0.010742 | 0.009164 | 0.007504 | 0.005954 | 0.004169 | 0.002368 | 0.001943 | 0 | -0.002654 | -0.005621 | -0.008258 | -0.010429 | -0.012436 | -0.013927 | -0.015568 | -0.01667 | -0.017421 | -0.017451 | -0.017091 | -0.016522 | -0.015442
0.044344 | 0.039083 | 0.035338 | 0.032505 | 0.030435 | 0.028666 | 0.026899 | 0.025444 | 0.023801 | 0.022115 | 0.020219 | 0.018305 | 0.016388 | 0.014467 | 0.012414 | 0.010708 | 0.009143 | 0.007486 | 0.005964 | 0.004124 | 0.002354 | 0.00196 | 0 | -0.00265 | -0.005618 | -0.008271 | -0.010468 | -0.012452 | -0.013889 | -0.015581 | -0.016684 | -0.017408 | -0.017447 | -0.017102 | -0.016567 | -0.015475
0.043698 | 0.038549 | 0.034873 | 0.03214 | 0.030111 | 0.028427 | 0.026659 | 0.025239 | 0.023643 | 0.021999 | 0.02009 | 0.01817 | 0.016285 | 0.014373 | 0.012326 | 0.010665 | 0.009127 | 0.007418 | 0.005869 | 0.004139 | 0.002335 | 0.001894 | 0 | -0.002636 | -0.005601 | -0.008212 | -0.010419 | -0.012369 | -0.01381 | -0.015464 | -0.016579 | -0.017301 | -0.017369 | -0.017006 | -0.016487 | -0.015446
0.042992 | 0.037952 | 0.03443 | 0.031771 | 0.029828 | 0.028173 | 0.026435 | 0.025009 | 0.023457 | 0.021807 | 0.019928 | 0.018035 | 0.016163 | 0.014258 | 0.012214 | 0.010569 | 0.00903 | 0.007387 | 0.005817 | 0.004059 | 0.002307 | 0.001876 | 0 | -0.002633 | -0.005548 | -0.008157 | -0.01031 | -0.012268 | -0.013705 | -0.015358 | -0.016455 | -0.01723 | -0.017285 | -0.016953 | -0.016432 | -0.015382
0.042221 | 0.037287 | 0.033844 | 0.031311 | 0.029386 | 0.02777 | 0.02605 | 0.024678 | 0.023098 | 0.02148 | 0.019673 | 0.017788 | 0.015942 | 0.014082 | 0.012065 | 0.01043 | 0.008902 | 0.007319 | 0.005796 | 0.004019 | 0.002336 | 0.001903 | 0 | -0.00258 | -0.005478 | -0.00805 | -0.01022 | -0.012159 | -0.013581 | -0.015243 | -0.016309 | -0.017066 | -0.017113 | -0.016791 | -0.016301 | -0.015272
0.042259 | 0.037349 | 0.033843 | 0.031212 | 0.029324 | 0.027686 | 0.025894 | 0.024543 | 0.023009 | 0.021379 | 0.019577 | 0.017679 | 0.015827 | 0.013948 | 0.011991 | 0.010371 | 0.008821 | 0.007181 | 0.005686 | 0.003942 | 0.002254 | 0.001815 | 0 | -0.002571 | -0.005427 | -0.007913 | -0.010022 | -0.011901 | -0.013293 | -0.014897 | -0.015993 | -0.016789 | -0.016829 | -0.016533 | -0.016052 | -0.015022
0.041813 | 0.036956 | 0.033568 | 0.030989 | 0.029042 | 0.027341 | 0.025648 | 0.024209 | 0.022695 | 0.021096 | 0.019255 | 0.017397 | 0.015595 | 0.013745 | 0.011793 | 0.010183 | 0.008689 | 0.007123 | 0.005605 | 0.003907 | 0.002182 | 0.001794 | 0 | -0.002475 | -0.005213 | -0.007699 | -0.009705 | -0.011536 | -0.012884 | -0.014471 | -0.015534 | -0.01628 | -0.016365 | -0.016077 | -0.01564 | -0.014598
0.041191 | 0.036487 | 0.03313 | 0.03058 | 0.028657 | 0.027019 | 0.025235 | 0.02387 | 0.022356 | 0.020745 | 0.018997 | 0.017106 | 0.0153 | 0.013555 | 0.011633 | 0.010029 | 0.008528 | 0.006951 | 0.005516 | 0.003756 | 0.00212 | 0.001737 | 0 | -0.002478 | -0.005183 | -0.007567 | -0.009516 | -0.011317 | -0.012641 | -0.014171 | -0.015294 | -0.016067 | -0.016119 | -0.015833 | -0.015372 | -0.01439
0.041316 | 0.036686 | 0.033323 | 0.030749 | 0.028815 | 0.027124 | 0.025339 | 0.023956 | 0.022436 | 0.020837 | 0.019099 | 0.017189 | 0.015386 | 0.013542 | 0.011703 | 0.01001 | 0.008595 | 0.006988 | 0.005557 | 0.003856 | 0.002173 | 0.001729 | 0 | -0.002371 | -0.005031 | -0.007387 | -0.009296 | -0.011088 | -0.012411 | -0.013935 | -0.015015 | -0.015812 | -0.015954 | -0.015646 | -0.015219 | -0.014239
0.040594 | 0.036156 | 0.032921 | 0.030345 | 0.028423 | 0.026787 | 0.024967 | 0.023568 | 0.022097 | 0.020502 | 0.018715 | 0.01688 | 0.015154 | 0.013349 | 0.011444 | 0.009842 | 0.008406 | 0.006887 | 0.005472 | 0.003834 | 0.002208 | 0.001762 | 0 | -0.002359 | -0.004949 | -0.007253 | -0.009128 | -0.010932 | -0.012257 | -0.013826 | -0.014916 | -0.015702 | -0.015837 | -0.015531 | -0.015132 | -0.014161
0.040902 | 0.036481 | 0.033261 | 0.030688 | 0.028658 | 0.026919 | 0.02513 | 0.02369 | 0.022177 | 0.020549 | 0.018804 | 0.016909 | 0.015159 | 0.013309 | 0.011517 | 0.009861 | 0.008476 | 0.006844 | 0.005416 | 0.003728 | 0.00215 | 0.001724 | 0 | -0.002349 | -0.005005 | -0.007269 | -0.009202 | -0.010964 | -0.012271 | -0.013833 | -0.014991 | -0.015757 | -0.015889 | -0.015651 | -0.015238 | -0.014273
0.040507 | 0.036244 | 0.03313 | 0.030466 | 0.02853 | 0.026796 | 0.024926 | 0.02356 | 0.022072 | 0.020436 | 0.018661 | 0.01686 | 0.015065 | 0.01327 | 0.011421 | 0.009779 | 0.00836 | 0.006812 | 0.005385 | 0.003716 | 0.002154 | 0.001667 | 0 | -0.002329 | -0.004917 | -0.00715 | -0.008986 | -0.010748 | -0.01212 | -0.013687 | -0.014838 | -0.01569 | -0.015805 | -0.015525 | -0.01519 | -0.01428
0.041044 | 0.036799 | 0.033687 | 0.031048 | 0.028919 | 0.027117 | 0.025318 | 0.023823 | 0.022242 | 0.020608 | 0.018821 | 0.016928 | 0.015111 | 0.013337 | 0.011433 | 0.009828 | 0.008398 | 0.006791 | 0.005387 | 0.003712 | 0.002119 | 0.001715 | 0 | -0.002324 | -0.004934 | -0.007156 | -0.009049 | -0.010837 | -0.012223 | -0.013854 | -0.014984 | -0.015731 | -0.015965 | -0.015671 | -0.015266 | -0.014353
0.040555 | 0.036599 | 0.033512 | 0.030801 | 0.028833 | 0.027067 | 0.025059 | 0.023686 | 0.022212 | 0.020591 | 0.018794 | 0.016943 | 0.015124 | 0.013358 | 0.011467 | 0.009842 | 0.008424 | 0.006759 | 0.005436 | 0.003789 | 0.002176 | 0.001741 | 0 | -0.002301 | -0.004878 | -0.007156 | -0.008984 | -0.010747 | -0.012094 | -0.013729 | -0.014934 | -0.015772 | -0.015945 | -0.015645 | -0.015365 | -0.014429
0.041255 | 0.037268 | 0.034189 | 0.031442 | 0.029348 | 0.027537 | 0.025548 | 0.024095 | 0.022551 | 0.02086 | 0.019111 | 0.01719 | 0.015338 | 0.013424 | 0.011584 | 0.009907 | 0.008457 | 0.006876 | 0.005451 | 0.003823 | 0.002184 | 0.001795 | 0 | -0.002366 | -0.004921 | -0.0072 | -0.009066 | -0.010916 | -0.012294 | -0.013876 | -0.015084 | -0.015942 | -0.016142 | -0.015939 | -0.015556 | -0.014637
0.040913 | 0.037028 | 0.034016 | 0.031371 | 0.029225 | 0.027473 | 0.025563 | 0.024032 | 0.022459 | 0.020756 | 0.019006 | 0.017181 | 0.015234 | 0.013434 | 0.011484 | 0.009981 | 0.008421 | 0.006786 | 0.005466 | 0.00377 | 0.002144 | 0.001753 | 0 | -0.002328 | -0.004935 | -0.007221 | -0.009137 | -0.010916 | -0.012381 | -0.014077 | -0.015156 | -0.016008 | -0.016194 | -0.015957 | -0.015613 | -0.014713
0.041393 | 0.037569 | 0.034407 | 0.031602 | 0.029597 | 0.02774 | 0.025714 | 0.024239 | 0.022792 | 0.021148 | 0.019347 | 0.017363 | 0.015581 | 0.013635 | 0.011807 | 0.01008 | 0.008627 | 0.006918 | 0.005511 | 0.003912 | 0.002226 | 0.00182 | 0 | -0.00241 | -0.005064 | -0.007395 | -0.00925 | -0.011126 | -0.01254 | -0.014145 | -0.015443 | -0.016283 | -0.016451 | -0.016201 | -0.015945 | -0.015044
0.040951 | 0.037312 | 0.034369 | 0.031704 | 0.029683 | 0.027905 | 0.025916 | 0.024412 | 0.022888 | 0.021281 | 0.019471 | 0.017552 | 0.015616 | 0.013696 | 0.011821 | 0.010135 | 0.008635 | 0.007039 | 0.005603 | 0.003912 | 0.002271 | 0.001834 | 0 | -0.002367 | -0.005058 | -0.00736 | -0.009276 | -0.011089 | -0.012576 | -0.014215 | -0.015469 | -0.016355 | -0.016579 | -0.016308 | -0.016053 | -0.015201
0.040935 | 0.037274 | 0.034211 | 0.031565 | 0.029575 | 0.027852 | 0.025841 | 0.024324 | 0.022907 | 0.021249 | 0.019518 | 0.017512 | 0.015696 | 0.013745 | 0.011892 | 0.010119 | 0.008716 | 0.007009 | 0.005669 | 0.003944 | 0.002395 | 0.00183 | 0 | -0.002496 | -0.005177 | -0.007554 | -0.009519 | -0.011402 | -0.01282 | -0.014484 | -0.015762 | -0.016545 | -0.016761 | -0.016595 | -0.016272 | -0.0154
0.040093 | 0.0366 | 0.033587 | 0.031006 | 0.02918 | 0.02754 | 0.025451 | 0.02404 | 0.022691 | 0.021072 | 0.019443 | 0.017463 | 0.015621 | 0.01376 | 0.011869 | 0.010155 | 0.008647 | 0.006947 | 0.005497 | 0.003935 | 0.00227 | 0.001769 | 0 | -0.002463 | -0.005114 | -0.007488 | -0.00939 | -0.011266 | -0.012804 | -0.014401 | -0.015692 | -0.016563 | -0.016851 | -0.016542 | -0.016291 | -0.015492
0.039625 | 0.036051 | 0.033269 | 0.030788 | 0.028898 | 0.027178 | 0.025397 | 0.023937 | 0.022455 | 0.020822 | 0.019138 | 0.017271 | 0.015433 | 0.013532 | 0.011717 | 0.00992 | 0.008499 | 0.006938 | 0.005516 | 0.003815 | 0.002354 | 0.001858 | 0 | -0.002427 | -0.005117 | -0.007479 | -0.009348 | -0.011248 | -0.012796 | -0.014439 | -0.015671 | -0.016483 | -0.016747 | -0.01653 | -0.016189 | -0.015389
0.038066 | 0.034632 | 0.031754 | 0.029361 | 0.027733 | 0.026199 | 0.024202 | 0.022922 | 0.02157 | 0.020153 | 0.018488 | 0.016771 | 0.014944 | 0.01318 | 0.011338 | 0.009672 | 0.00832 | 0.006608 | 0.005312 | 0.003799 | 0.002156 | 0.001704 | 0 | -0.002334 | -0.004932 | -0.00724 | -0.009094 | -0.010882 | -0.012353 | -0.01393 | -0.015171 | -0.016008 | -0.016249 | -0.016054 | -0.015826 | -0.014987
0.038646 | 0.034949 | 0.031811 | 0.029305 | 0.027519 | 0.025948 | 0.023901 | 0.022599 | 0.021445 | 0.019893 | 0.018304 | 0.016449 | 0.014839 | 0.012967 | 0.011212 | 0.009636 | 0.008164 | 0.006431 | 0.005248 | 0.003735 | 0.002132 | 0.001708 | 0 | -0.002305 | -0.004751 | -0.006966 | -0.008634 | -0.010454 | -0.011866 | -0.013413 | -0.014739 | -0.015611 | -0.015817 | -0.015515 | -0.015443 | -0.01454